## Endoscopy Tower
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Endoscopy Tower System With Endoscopes | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
System must state of the art , latest release model , having Egonomic design and easy to use
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
Specify Unit Make ,Model and country of origin
1 | Mobile Endoscopy Trolley
1.1 | Equipped with a monitor's arm having both swivel and tilt functionality
1.2 | Equipped with scopes' holding pole
1.3 | Equipped with a cable management system
1.4 | Equipped with a suction machine
1.5 | Equipped with an irrigation/flushing pump with padel and tubing
1.6 | Equipped with all standard and optional accessories
2 | Video System Center
2.1 | Video System with LED Spectrum Light
2.2 | 4K Resolution output
2.3 | It can be connected to many different types of endoscopes. i.e : Gastroscope,Colonoscope, Duodenoscope, Bronchoscope….
2.4 | Having conventional white light imaging
2.5 | Having Narrow Band Imaging (NBI)
2.6 | Having Auto Fluorescence Imaging (AFI)
2.7 | Having Texture & Color Enhancement Imaging (TXI)
2.8 | Having Red Dichromatic Imaging (RDI)
2.9 | Having brightness adjustment with maintenance of contrast (BI-MAC)
3 | Having Extended Depth of Field (EDOF)
3.1 | Having Automatic gain control (AGC)
3.2 | Having pre-freeze function
3.3 | Having image size selection
3.4 | Having electric zoom
3.5 | Having Automatic white balance
3.6 | Having analog and digital signal outputs
3.7 | Can store several users' functions settings
3.8 | Having remote control
4 | 4K LCD Monitor (32-inch approximately)
5 | Videoscopes should be of latest technology
5.1 | Adult Video Gastroscope :
a. | Regular/Routine adult gastroscope | 1
b. | Single channel therapeutic gastroscope | 1
5.2 | Adult Video Colonoscope :
a. | Regular/Routine adult colonoscope with variable/adjustable stiffness | 2
5 | Therapeutic Dudonoscope (ERCP) | 1
6 | Portable Suction must be included
7 | Endoscopy monitor must be able to connect slave wall mounted monitor for better viewing of the surgeon and installtion must be included in the offer .
8 | Unit must be provided along with image and video recording system , which is also having HL7 connectivity and can also connected with HIS system.
9 | All scopes must be supplied along with disinfection adpaters which are compatable with our Madivator-DSD 201
10 | Unit must be antimicrobial Protection
11 | Specify Dimensions and weight of the units
12 | Unit should be provided with all necessary accessories and consumables.
13 | Must have universal power adapter 220-240v ,50/60HZ
14 | Specify the power consumption for the offred units and heat dissipation(BTU/hr)
1 | Quality and Patient safety
Unit must be FDA / CE certified
Unit must comply ISO 13485 Quality management system for medical device or equivalent .
Unit must comply with Eletrical safety test IEC 60601-1,-2,-18 or equivalent
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and End user / application training must be done for Endoscopy doctors and Nurses.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warranty .

## Endoscopy drying cabinet
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| ENDOSCOPE DRYING CABINET-Store up to 10-12 full size endoscopes | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | System must state of the art , latest release model , having Egonomic design and easy to use.
2 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
3 | Specify Unit Make ,Model and country of origin
4 | Innovative Designs and High-Quality Construction
5 | Available in a range of sizes and features to fit your endoscope storage needs
6 | Non-porous chamber construction using stainless steel
7 | Clear visibility through tempered glass doors and internal lighting and UV system synchronous with door,
8 | Gasket-less doors and password or clock protected doors
9 | scope arrangement -Vertical Hanging Racks
10 | Scope holders with pull-out hanging rails
11 | Capacity Options up to 10-12 channels
12 | Bottom and middle parts adjustable suspended holding apparatuses,
13 | Reusable drying connectors, including auxiliary water
14 | HEPA-Filtered Air
15 | Add-On Options for an additional scope ?
16 | Channel Purging Connectors
17 | Scope Hang Time Tracking with Touch Interface and Barcode Reader
18 | specify Dual filtered air maintains positive cabinet pressure
19 | Temperature and humidity monitoring within the cabinet
20 | Advanced fully Independent Monitoring System (IMS)
21 | HEPA filtered and dehumidified air delivered to each scope channel
22 | Must have guaranteed dry scopes within 3-5 hours
23 | High flow rate channel drying with 6 ___LPM per scope
24 | Must have Temperature and humidity monitoring within the cabinet
25 | Dual filtered air maintains positive cabinet pressure
26 | Clear toughened glass doors allow easy scope recognition
27 | Secure mounting for scope control section and lightguide plug
28 | Double door lockable front cover,
29 | Air Delivery System or Hospital Air integration sold separately
30 | Dry Air compressure with moisure free and oil free
31 | Noise less compressor
32 | Air manifold with a quick disconnect
33 | Anti bacterial and filtered air discharges that dries internal channels of scope
34 | Silent and non-vibrating fan system,
35 | Efficient Storage in a Controlled Environment
36 | Storage – Accommodate a wide variety of scope sizes, and vertical hanging racks prevent scopes from touching and damaging each other
37 | Controlled Environment – Air exchange every___ seconds on average, while filtration removes 99.97% of particulates1
38 | Internal temperature of cabinet can be fixed between 25ºC – 40ºC,
39 | UV system that operates throughout the most effective wave (254 NM),
40 | User friendly safety systems (UV does not operate when door is open),
41 | 6 mm. thick tempered glass door that prevent transmission of ultraviolet light,
42 | Units muist offers full traceability of scopes -Internal Data Archiving
43 | Tracking – Available inventory can be viewed at a glance, and swipe badge access maintains scope security
44 | Unit must have printer facility
45 | specify touchscreen size with individual pin-code protection
46 | Interface that can be controlled through LCD graphical screen,
42 | Unit must be antimicrobial Protection
44 | Internal & Externa lDimensionss in cm/Inch
44 | Unit should be provided with all necessary accessories and consumables.
44 | Must have universal power adapter 220-240v ,50/60HZ
44 | Specify the power consumption for the offred units and heat dissipation(BTU/hr)
45 | To povide the channels adaters for Below metioned scopes
| Name | Model | Manufacturer
1 | VIDEO BRONCHOSCOPE | BF-6C260 | OLYMPUS, JAPAN
2 | VIDEO BRONCHOSCOPE | BF-F260 | OLYMPUS, JAPAN
3 | SIGMOIDOSCOPE | PCF-240S | OLYMPUS, JAPAN
4 | VIDEO COLONOSCOPES | CF-H260AL | OLYMPUS, JAPAN
5 | VIDEO COLONOSCOPES | PCF-Q260AL | OLYMPUS, JAPAN
6 | VIDEO DUODENOSCOPE | TJF-260V | OLYMPUS, JAPAN
7 | VIDEO BRONCHOSCOPE | BF-XP260F | OLYMPUS, JAPAN
8 | VIDEO GASTROSCOPE | GIF-H260Z | OLYMPUS, JAPAN
9 | VIDEO COLONOSCOPE | CF-H260DL | OLYMPUS, JAPAN
10 | VIDEO BRONCHOSCOPE | EB19-J10 | PENTEX
46 | Quality and Safety Standards
Offered Unit must be FDA / CE marked .
Unit must comply with BS EN 16442 and ISO 13485 Quality management system for medical device or equivalent .
System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1
47 | Accessories
Supplied with all standard accessories.
48 | Consumables
Supplied with all standard consumables.
49 | Avialable Additional Option
OPTIONS DESCRIPTION
50 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and for Neuro and EEG Team (End User Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warranty .