## Mobile X-ray
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Digital Mobile X-Ray unit | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
Mobile DR X-ray Machine with sliding column, Maximum 30-40 kW ,500 mA,17x17,10x12 Wireless detectors with high image resolution up to 4 Lp/mm , Nano model with fast image transmission capacity
2 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
3 | Integration with PACS , GE Dose Watch & RIS :Unit must be able to integrate with hospital PACS system & GE Dose Watch .All softwares and hardwares must be provided by vendor to do the intergation .
3.1 | A fully digital radiography system capable of detector exposure in vertical, horizontal and oblique positions to perform general radiography. The unit should be completely integrated (integrated Generator and Image Acquisition) and comprise the following along with auto quality control features incorporated.
4 | X-ray Tube and Generator
4.1 | X-ray tube with high anode heat capacity & cooling rate (Min250W)
4.2 | Specify Heat storage capacity of anode Minimum(Heat Storage: 30000J)
4.3 | Power rating, kW @ 100 kVp
4.4 | X-ray power ≥ 80 kW
4.5 | Tube voltage:40-140 Kv with 1.0kv increment
4.6 | mAs range:0.1 – 500 mAs, 42selections
4.7 | Exposure times :0.001-1.25 s
4.8 | Focal spot (S/L):0.3/1.0
4.9 | PACS & DICOM interface for Image store, Query/Retrieve, Print, Work List Generation, MPPS, and archive
4.1 | mA Range: 30~130mA, variable
4.11 | Console & Display: kV & mA & mAs/sec digital display Density (for human) digital display Prep. & Exposure LED Collimator lamp Switch APR Store Switch
4.12 | Type: Fixed anode Focal spot size: 1.8mm Inherent Filtration: 0.5mm Al Target angel: 15° Collimator: Manual: PCM100 Lamp: 24V150watt, halogen lamp, built-in: 30sec. Timer (DC: 30sec) X-ray Field Size: 35cm * 35cm @ 65cm SID - Charge input: 100V – 240 VAC, 50/60Hz
5 | Tube Movement
5.1 | Horizontal, cm (in)
5.2 | Vertical, cm (in)
5.3 | Rotation
5.4 | Z-axis, °
5.5 | X-axis, °
5.6 | Aluminium Filter mm
6 | Wireless portable detector
6.1 | Type :CsI(Scintillator Cesium iodide)
6.2 | Housing:Carbon fiber
6.3 | Detector sizes:17 x 17 cm; 10 x 12 cm
6.4 | High Image resolution up to 4 Lp/mm
6.5 | Secify weight of each FPD
6.6 | Pixel size:148 µm
6.7 | Matrix Depth 16 bit/Pixel and matrices of 4288×4288
6.8 | FPDs must be IPX7 or equlivent water proof standard
6.9 | Specify number of physical drop allowed for FPDs
6.1 | wireless router
6.11 | The resolution should be minimum of 3.0 lines pair/millimeter. - The pixel size should be 130 um of less. - Detector Quantum Efficiency (D.Q.E) should be more 60% @ Zero Line pairs. - The active matrix size should be 3k x 3k or more.
6.12 | 17.5 MegaPixel sensor in terms of the number of Dedector Eliments
7 | Type of power source
7.1 | AC line
7.2 | Battery type
7.3 | Separate batteries for drive and generator control
7.4 | Energy capacity, mAs @ 80 kV?
7.5 | power of charging -Must have single phase 230V/50-60Hz
7.6 | Charging time, full charge, hr?
7.7 | Can also make exposures using on AC power?
7.8 | Number of images per battery charge
7.9 | Time to display image, sec
7.1 | LineVoltage Compensator
7.11 | Protection level against electric shocks: type B
7.12 | Error code display: - Tube overheating - High voltage, power supply not ready - Exposure request which exceeds the max power the generator exposure request for too many mAs for the exposure - The exposure is impossible due to the absence of the emission characteristics of the tube
8 | Clinical workflow
8.1 | The system must have grate forward visibility and easy maneuvers to support comfortable operation
8.2 | Must have large touch panel display for quickly viewing images
8.3 | specify system start up time
8.4 | Must have Large Image Storage Capacity
8.5 | Shock-resistant body design
8.6 | Fast and easy positioning
8.7 | system access Numeric keypad / PIN code
8.8 | Upgradeable to higher version
8.9 | footprint of the machine (LxBxH)
8.1 | User friendly operator console
8.11 | Hand exposure control
8.12 | SID RANGE, cm (in)
8.13 | Integrated DAP Monitor with shortest warm-up time
8.14 | Focal point distance from floor 55 to 205cm:
9 | Unit must be antimicrobial Protection
10 | Specify Weight and Dimensions of the unit
11 | Power consumption & Heat dissipation in BTU/h
12 | High Imaging Resolution Unit for Medical Mobile Application
13 | Quality control ,quality assurance and accaeptance test should be done by the company
14 | Safety assessment should be done by the company
15 | Unit must be available on demo basis
16 | Accessories
16.1 | All necessary Accessories, consumables and disposables for starting-up.
16.2 | General Radiology QC tools for performance verification done by Enduser, if required.
16.3 | Lead aprons Quantity ( double face 0.35 mm lead equivalent Adult 2 and Light weight paediatric use 1/2 apron 3 Numbers
16.4 | Thyroid shield protection Quantity 3
16.5 | DICOM Image Storage Service Class as defined below: The DICOM Storage Commitment Service Class
16.6 | For retrieving prior images from the PACS, the DICOM Query/Retrieve (C-FIND and C-MOVE for Patient Root and/or Study Root Model) Service Class should be support.
17 | Quality and Patient safety
17.1 | Unit must be FDA / CE certified
17.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
17.3 | Unit must comply with Eletrical safety test IEC 60601-1,-2,-18 or equivalent
17.4 | User-friendly operation and smooth movement - One-unit package with in-built battery/generator system - Durable with operational implication for different position and parts of the body. - Net weight not more than 110 kg.
18 | Accessories
18.1 | Supplied with all standard accessories.
19 | Consumables
19.1 | Supplied with all standard consumables.
20 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for OT Anesthesia technicians and Anesthesia doctors (End User Training) is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Continuous Update and Upgrading of the system and all provided software with the system including and limited to DICOM software, during the warranty (any hardware upgrade needed for such software upgrade should be included and provided with no additional cost by the vendor). It is expected from the vendor to always install the latest software versions within 4 months of its official release.
9 | Maximum time needed for parts ordering and replacement (is calculated as a system down time )
10 | Maximum response time for system faults or breakdown calls (max. time for the engineer to be at the site) to be provided.
11 | Warranty must be at laest 3 years comprehensive from date of installation and equipment acceptance.