## MRI Machine
Indicative Specifications / Evaluation Form
Item No. | Item Description | To Be Filled By Vendor |  | To be filled
| Open MRI Machine ≥1.0 Tesla | Comply | Supporting Documents Ref. | by Hospital
| Specifications | Yes/No (If Yes Specify)
State of Art High Field whole body Open MRI system,Magnetic field strength ≥1.0 Tesla with active shielding and External Interference Shielding ,with wide bore ≥45cm,Maximum available channels ≥ 32 is preferred with superconducting magnet and latest digital technology ,low noise, fast imaging, enhanced patient experience, having all the anatomy specific coils and software for all types of application including and not limited to Neurology, Paediatrics, Oncology, Musculoskeletal, Abdomen, liver, Body, Prostate, Breast,Vascular, Functional procedures,cardiac .. etc. all features expected from each of these application to be offered as standard, other avialble applications,upgradable features should be clearly mentioned as optional.
1 | Installation and B₀ main field
1.1 | All pre-installation requirements should be submitted along with the quotation including suggested room dimensions etc.
Power Supply Capacity and Specifications
Minimum Space Requirement for your system
Minimum installed area (m2)
Faraday cage(shild) included, Technical room, UPS, chillers
All room preparation and system installation requirements (to be clarified and agreed on after vendor initial site visit)- full room layout required.
Installation should be done by original manufacturer trained Engineer
1.2 | Scanner FDA nod
1.3 | Scanner CE certification
1.4 | Field strength:
1.5 | Dimensions of magnet bore
Patient Aperture Width cm at narrowest: (width x height)
Magnet bore length (w/o covers):
1.6 | Fringe field
Is the magnet equipped with magnetic shielding?
Distance of the 0.5-mT line from isocenter
horizontal axis (x):
vertical axis (y):
longitudinal axis (z):
1.7 | Shielding against external interferences
Is the magnet shielded against external interferences? (e.g. moving ferromagnetic objects, transformers)
Is shielding against external interferences active during a scan?
Description:
1.8 | Magnet coils
Number of magnet coils:
Are magnet coils actively shielded?
1.9 | Main magnetic field Homogeneity of <0.5 ppm over a 35 cm DSV and Automatic improve field homogeneity
Max Field of view in cm
Inside a volume of cm³
1.10 | Long term stability of the magnetic field:
1.11 | Helium Consumption
| Zero helium boil-off technology | yes/No with details
Helium refill frequency
Tank capacity of liquid helium at 100% level:
1.12 | Shim System for the Installation
Passive Shim for improvement of the field homogeneity (e.g. with iron):
Active Shim for improvement of the field homogeneity(Active shimming preferred) and , dynamic FOV shimming
Number of independent shim terms with the active shim (e.g. 8, for the 3 linear terms and for the 5 nonlinear terms):
1.13 | Patient-specific Shim
Is active shimming for a patient-specific improvement of the field homogeneity possible?
Number of independent shim channels with the active shim (e.g. 3 for the 3 linear channels):
Fully automated patient-specific shimming:
Typical duration of the automatic shim:
1.14 | Advanced shim system for patient-specific shim
Active shim system (with non-linear shim terms) for a patient-specific improvement of the field homogeneity:
Number of independent shim channels with the advanced active shim system (e.g. 3 for the 3 linear channels):
Fully automated patient-specific shimming:
Typical duration of the automatic shim:
1.15 | RF CAGE & STANDARD ACCESSORIES
RF cage - RF room with one Door and one View window (min 140 x 100) - state full technical specifications of the offered RF cage.
The maximum acoustic noise level expected outside the MRI room (state the technical specification of the door of MRI room offered - with its sound (acoustic noise) proof
MR Room (zone 4) and zone 3 Oxygen Monitors
Door Mounted Metal Detector
Safety marking plate with Alarming System in case of Chiller/ Magnet/Power down.
2 Quench buttons (in zone 4 and above MR console)
Measuring and Adjustment Kit : Includes valves, Flow meter, water filter, shunt valve & fittings
1.16 | In Bore patient comfort, and monitoring features
Feet-First or head -first entry
In-bore lighting system
Audiovisual device for patient comfort and funtional research
In-bore patient ventilation system
Patient accessories: Dedicated head arm support, Table mattress set, Head/leg support, Knee support, Positioning wedges, Small foam wedges and pads, Set of sandbags, positioning pads, Immobilization straps, Set of patient fixation straps etc
Self-storing non-ferrous IV pole
Comprehensive set of pediatric accessories : Includes Baby support mattress for small infants, Anterior coil frame etc.
2 | Gradient System
2.1 | The required gradient performance has a peak gradient field strength ≥32 mT/m per axis and a slew rate ≥100 T/m/s per axis simultaneously or Equivalent slew rate (in x, y, z-direction)
For gradient systems with several operational modes, the specifications for each gradient mode have to be provided.
2.1.1 | Describe the shape of the gradient linearity volume
2.1.2 | Are the gradient coils actively shielded?
2.1.3 | Maximum gradient amplitude
2.1.4 | Maximum slew rate
2.1.5 | Minimum rise time from 0 to max. gradient amplitude:
2.1.6 | Duty cycle:
2.1.7 | Manufacturer of the gradient coil: Duty-Cycle 100%, full gradient performance for long term scanning
2.1.8 | Maximum output current per axis of the Gradient Power Amplifier (GPA) horizontal axis (x, "left-right"), vertical axis (y, "anterior-posterior"), longitudinal axis (z, "head-feet")
2.1.9 | Maximum output voltage per axis of the Gradient Power Amplifier
2.1.10 | horizontal axis (x, "left-right"), vertical axis (y, "anterior-posterior"), longitudinal axis (z, "head-feet")
2.1.11 | Maximum output voltage and output current available in one single gradient pulse
2.1.12 | Manufacturer and type of the gradient amplifier:
2.1.13 | Maximum Field-of-View for the above-mentioned gradient specifications (depending on application) horizontal axis (x, "left-right"), vertical axis (y, "anterior-posterior"), longitudinal axis (z, "head-feet")
2.1.14 | Minimum TR ms Spin echo
2.1.15 | Minimum TR ms 2D gradient echo
2.1.16 | Minimum TR ms 3D gradient echo
2.1.17 | Gradient performance will be rated according to its clinical impact, like resolution, TSE factor, 2D/3D slice thickness, and EPI parameters give Details
2.1.18 | Noise reduction for different scans mode should be included, mentioned and will be rated
3 | Radio Frequency (RF) Transmitter System B1 field
3.1 | Signal generation and processing, RF technology?
3.2 | Frequency resolution of the RF synthesizer:
3.3 | Bandwidth of the RF transmitter:
3.4 | Max. power of the transmitter amplifier ≥18kW transmit power at 100% efficiency
3.5 | RF Multi Transmit Capability
3.6 | Is the amplitude and phase transmission optimized for dedicated body regions
3.7 | Is the transmission system integrated in the magnet housing?
3.8 | 1.5 MHz Receiver sampling rate per each channel.
3.9 | Coils integration is mandatory point of evaluation “more than2 coils simultaneously to be connected especially for Liver and abdomen imaging” additional to different applications over the body to show the maximum number of coil elements over defined FOV.
4 | RF Receiver System
4.1 | Are the receive components integrated into the magnet housing?
4.2 | Maximum number of channels. Minimum of 16 channels, anything higher is preferred
Signal receiving and processing
4.2.1 | Number of independent receiver channels that can be used simultaneously in one single scan and in one single FOV, each generating an independent partial image
Partial Images to proof independent channels to be provided for at least 5 body regions (e.g. Head, Head/neck, Body, ...)
4.2.2 | Max. receiver bandwidth of each receiver channel:
Do all channels have the same bandwidth?
4.2.3 | Sampling rate of each ADC:
4.2.4 | Receiver signal resolution:
4.2.5 | Noise of the preamplifier:
4.3 | Max. number of Coil Elements
4.3.1 | Max,number of available coils that can be used in FOV Designs to enhance SNR as will as parallel image compatibility should be offered (for every coil, parallel Imaging reduction factor should be stated), Coils, Higher Channels / Wider coverage will be preferable and Automatic Coil Switching
4.2.6 | Max. number of simultaneously connected coil elements:
4.3.2 | Can all of these coil elements be seamlessly used in one study without coil changes and without repositioning of the patient?
4.3.3 | Provide an example of the max. number of coil elements that can simultaneously be connected and used in one examination without coil or patient repositioning:
4.3.4 | Body coil for transmission & receiving.
5 | RF Coils
Special volume and surface coils for the following applications shall be available with the system. The required coils have to be best in the market with the maximum number of elements and to be included in the quote.
5.1 | Applications
5.1.1 | Multi-channel array coils for high-speed Parallel Acquisition Techniques:
5.1.2 | Head
5.1.3 | Head-Neck combined, Neuro-Vascular
5.1.4 | Cervical Spine, Thoracic Spine, Lumbar Spine, Whole Spine CTL array coil or equivalent for imaging whole CNS without patient repositioning
5.1.5 | Whole CNS (Head + Whole Spine):
5.1.6 | Chest, Heart:
5.1.7 | Abdomen: High Resolution 3D, MRCP:
5.1.8 | Pelvis and Hips:
5.1.9 | Whole Torso:
5.1.10 | Short cable and light-weight coils, array coils
5.1.11 | Knee, Ankle and Foot:standard
5.1.12 | Shoulder:Rigid coil as standard with Large flex coil
5.1.13 | Elbow:Rigid coil as standard with Large flex coil
5.1.14 | Wrist:Rigid coil as standard with small flex coil
5.1.15 | Breast:Rigid coil as standard with medium flex coil
5.1.16 | Whole-Body scanning with high-SNR surface coils:
5.1.17 | Which high-SNR surface coils and/or coil combinations can be used simultaneously for whole-body scanning?
5.1.18 | Dedicated XXL Vascular coil for Peripheral Angiography FULL coverage of lower limbs
5.1.19 | The ability for multiple coils connections and integrations at different procedures, No. of channels as well as parallel Imaging reduction factor should be stated for coils
5.1.20 | Gradient Motion Rephasing (GMR) Post-Processing, MIP, Surface Rendering Single Voxel and Multi voxel MR Spectroscopy Prostate Spectroscopy packagePerfusion imaging with all available maps Motion Correction techniques should be offered to overcome the motion in different directions at different contrast
5.1.21 | Unique features with clinical impact and additional software will be rated.
5.2 | Coil Technology
5.2.1 | Is it possible to connect multiple coils simultaneously?
5.2.2 | How many different coils can be connected and used simultaneously?
5.2.3 | Is it possible to combine coil elements from different coils, i.e. use coil elements from different coils simultaneously in one single scan?
5.2.4 | May certain coils stay on the patient table for the next examinations, in order to minimize coil changes?
5.2.5 | Please list these coils and name possible limitations:
5.2.6 | Are connected surface coils automatically detected by the system?
5.2.7 | Does this apply to all surface coils?
5.2.8 | Is the position of the coils in relation to the patient's anatomy automatically detected by the system and visualized in the User Interface?
5.2.9 | Does this apply to all surface coils?
5.2.10 | Is it visualized at the scanner what coils are plugged in?
5.2.11 | What kind of coil plugging mechanism is available?
5.2.12 | Are coils available without cables?
5.2.13 | Please list these coils:
5.2.14 | If cables are used, is the plugging mechanism operable with one hand?
5.2.15 | If cables are used, are the coil cables connected to the patient table instead of the magnet front, to ensure easier handling with shorter cable lengths?
5.2.16 | Kinematic coil and software(functional joint study)
5.2.17 | Can the patient table be moved freely (horizontally and vertically) with the coil cables plugged in?
6 | Patient Handling
6.1 | Patient Table Fixed or Standard
6.1.1 | Specify Vertical table movement and Max. vertical speed
| Minimum table hight from floor | cm
| Maximum table hight | cm
6.1.2 | Horizontal table movement
| Maximum horizontal range of table movement: | cm
| Maximum speed: | cm/s
| Accuracy of repositioning: | mm
6.1.3 | Maximum patient weight ≥300 kgs (including vertical and horizontal table movements): | kg / lbs
6.1.4 | Table Operation/ Landmark patient
| Accuracy of repositioning , +/- 1 mm, with Digital display of table position: | yes/no
| Isocenter positioning just by push-button: | yes/no
| Home positioning (out) just by push-button: | yes/no
| Home positioning (up) just by push-button: | yes/no
| Automatic positioning of patient table w/o laser: | yes/no
| Manual table move in case of emergency: | yes/no
| Table control from console (e.g. for peripheral MRA): | yes/no
6.3 | Scannable range / whole body coverage
| Is it possible to start and stop a scan at the magnet control panel? | yes/no
| Integrated infusion stand | yes/no
6.4 | In-room information for patient preparation
| In-room display of patient’s name | yes/no
| In-room information on the exam | yes/no
| In-room display of coils being plug in | yes/no
6.5 | Accessories for Patient Positioning
6.6 | All necessary accessories for patient positioning (mattresses, cushions, vacuum cushions) shall be included.
6.7 | Patient Comfort, Active patient alarm, Music entertainment for the patient (Integrated, incl. CD/-disc/Blutooth interface), intercom system.
7 | Physiological Signals
7.1 | Physiologic Measurement Unit (PMU)
All the trigging techniques should be standard. The following lead and probe need to be included as standard: - One set of patient wires, - One box of adult and pediatric size electrodes and - One fiber optic peripheral gating probe
7.1.1 | ECG
Is the ECG signal acquired from three ECG channels, to ensure optimal quality? (Vector Cardiogram, VCG)
7.1.2 | Peripheral Pulse Triggering
7.1.3 | Respiratory Triggering
Is respiratory triggering possible?
Pneumatic cushion, to be placed on the chest, for easy handling?
If a belt needs to be wrapped around the patient, please provide details.
7.1.4 | External Triggering
Is an interface included, for a trigger input from external sources?
Optical trigger output?
7.1.5 | Handling
Wireless trigger signal transmission for easier placement of the trigger devices:
State trigger devices that operate wirelessly(e.g. ECG, peripheral pulse, respiratory)
Are ECG, respiration, and SPO2 pulse displayed at the main console?
Are the trigger points visualized on the display for better control?
7.2 | Is an additional display of the physiological signals in the examination room available?
Applications
8 | Neuro Imaging
8.1 | Neuro Imaging Applications
8.1.1 | Comprehensive routine neuro imaging shall be included with SWI, CISS, Perfusion, spectroscopy and FiberTrak advanced imaging and processing methods for assessment of white matter fiber tracts in the brain).
Routine T1-weighted, T2-weighted,Stir and proton density weighted imaging for head and spine imaging:
Fluid-attenuated inversion recovery (FLAIR, 2D and 3D) imaging for complete suppression of the CSF signal:
SUSCEPTIBLE WEIGHTED IMAGING (SWI)
3D NERVE VIEW
Single-shot myelographic imaging with high resolution (true 512 matrix):
Multi-step imaging of the whole CNS (head + whole spine) with surface coil image quality, without repositioning of the patient or coils:
Multi-step imaging of the whole CNS (head + whole spine) with fully automated, software-controlled table movement:
8.1.2 | Diffusion Imaging
Diffusion imaging with multi shot EPI, HASTE and FSE and a high-resolution (non-single-shot) technique shall be included.
High-resolution diffusion imaging (non-single-shot):
Spine diffusion imaging for differentiation between osteoporosis, tumor infiltration and radiotherapy-induced changes:
Max. b-value:
Sequence should be able to carry out slice-by-slice tuning of resonance frequency, transmitter voltage, first order B0-shim and B1-shim. For whole-body diffusion, the SliceAdjust technology helps to avoid station boundaries and apparent broken spine artifacts as well as to preserve the SNR for whole-body diffusion.
Calculation of ADC (Apparent Diffusion Coefficient) maps:
8.1.3 | Diffusion Real-Time Evaluation
A fully automated real-time evaluation package for the evaluation of diffusion-weighted images (DWI) shall be included.
Can ADC maps (Apparent Diffusion Coefficient) be calculated automatically in real-time from DWI?
Can Trace-weighted images be calculated automatically in real-time from 3 separate orthogonal DWI scans with minimal TE?
8.2 | Neuro Workflow Features
Is the head of the patient automatically placed at the isocenter without any laser marking required?
Does the software guide the user step-by-step through the standard brain examination by showing example images and given explanation in text form?
How many mouse clicks does the user need to adapt an examination to the patient’s condition, e.g. switch to a protocol with motion correction if the patient is moving?
Are the protocols automatically positioned and aligned according to the anatomically derived sagittal, coronal, and axial slices of the localizer?
Are clinical decision steps possible so that the user can add or eliminate protocols or groups of protocols with one mouse click?
How many mouse clicks does the user need to rerun a sequence with identical parameters?
9. Angiography
9.1 | MR Angiography and Venography Applications Techniques for optimized vessel contrast and venous Suppression.
9.1.1 | Time-of-Flight MRA (TOF)
MR-Angiography using the Time-of-Flight method shall be included.
2D TOF:
3D TOF:
Sequential 2D TOF with "travelling saturation bands":
3D multi-slab (several overlapping 3D slabs in a single scan for larger anatomical coverage):
Automatic elimination of overlapping slices for 3D multi-slab:
Ramped RF-pulses for 3D scans, in order to decrease saturation effects in distal vessels:
9.1.2 | Phase Contrast MRA (PC)
MR Angiography using the Phase Contrast method shall be included.
9.1.3 | Contrast-enhanced MRA (ceMRA)
Ultrafast sequences for contrast-enhanced MR Angiography shall be included. Vascular Application, (all regular Vascular imaging sequences and the advanced application need to be included as standard), all available technique should be offered as a standard or optional.
9.1.4 | Bolus Timing
A software package for online-visualization of a fast 2D bolus scout and a fast switch to the 3D ceMRA scan shall be included.
9.1.5 | Neuro Functional Procedures
10 | Cardiac Morphology
10.1 | Cardiac Morphology Applications
10.1.1 | Is a step-by-step procedure provided for fast cardiac view creation (all relevant scouts)?
Graphical ‘3-point planning' of cardiac views:
Dark-blood preparation pulses, for optimized myocard-blood contrast:
Breath-hold techniques with high contrast:
10.1.3 | Cardiac Function
The assessment of ventricular function and wall motion shall be included.
Spoiled GRE technique (FLASH, Spoiled GRASS, T1-FFE) for cardiac function:
Gradient Echo with full transverse rephasing (TrueFISP, Balanced FFE, FIESTA) for cardiac function:
Automatic adjustment of the acquisition window to the current heart rate:
Graphical display of data acquisition window (e.g. TR) with respect to the ECG signal:
Can all these techniques be performed during breath-hold in order to examine the whole heart in a short time without artifacts?
Parallel Imaging (T-PAT, iPAT, SENSE, ASSET) for highest temporal and spatial resolution:
Radial scanning for real-time cine studies:
Real-time cine TrueFISP/FIESTA/B-FFE for non-triggered, free breathing acquisitions:
Is echo sharing and retrogating supported?
Are stress/rest examinations supported?
MyocardiaL Tagging
Look Locker for optimum Inversion Delay
Provides single breathhold, multi-echo, ECG-triggered acquisitions to provide T2*, R2*, T2 and R2 maps for assessment of myocardial tissue characteristics
Are stress/rest examinations supported?
Can arbitrarily oriented slices (short axis AND long axis simultaneously) per heart beat be acquired?
TI scout for optimization of contrast between infarction and normal myocardium:
2D and 3D 'delayed enhancement' techniques with segmented IR TrueFISP/FLASH sequences, for the assessment of myocardial viability:
2D 'delayed enhancement' techniques with segmented IR TrueFISP/FLASH sequences, for the assessment of myocardial viability:
First pass myocardial perfusion imaging with multiple slices per heart beat on the basis of fast GRE or SSFP techniques?
Phase-sensitive IR TrueFISP/FLASH for “Delayed-Enhancement“ studies with optimized contrast; no need to adjust the inversion time TI?
10.1.8 | Flow Quantification
A comprehensive package for flow quantification shall be offered as an option.
ECG-triggered phase contrast technique:
Segmented 2D GRE with echo-sharing, to increase the temporal resolution:
Breath-hold techniques for flow quantification:
Retrospective reconstruction algorithms for full R-R coverage in combination with arrhythmia rejection:
10.1.9 | Liver Imaging
Liver imaging application: triggering techniques and new strategies for motion control and fast imaging including MRCP
Quantitative MR for hepatic diseases assessment
Mention all available sequences and Motion Correction for Abdominal Imaging.
MR elastography
Dynamic and AI Imaging Sequences
10.1.10 | Oncology Application
In addition to all regular imaging sequences, the following advanced application need to be included as standard:
Whole Body screening with T1 sagittal, T1 Coronal, T2 Coronal, STIR and Diffusion Weighted Imaging, DWIBS
Whole body imaging using integrated body coil along with automated table stepping
11 | Breast Imaging
Breast Applications
11.1 | Suitable techniques to perform MR Breast Imaging and evaluation shall be included, in combination with a dedicated breast array coil.
11.2 | Availability of fast 3D Gradient Echo sequences for dynamic contrast studies (e.g. 64 slices with 256² matrix in < 1 minute):
11.3 | T2-weighted imaging with fat saturation (examination of cysts):
11.4 | 'Silicon-only' sequences for the examination of implants:
11.5 | Fully automated calculation of parametric maps for dynamic breast examinations (e.g. TTP, PEI):
11.6 | Breast coil compatible with Parallel Imaging:
11.7 | Diffusion weighted imaging with ADC mapping at higher b values
12 | Orthopedic Imaging
12.1 | Orthopedic imaging shall be included with joints cartilage assessment
Metal artifact reduction
mDixon Software
13 | Whole-Body Imaging
13.1 | Patient table for Whole-Body Imaging
Does the table shift range allow for whole-body examinations without patient re-positioning?
13.1.1 | Can whole-body examinations be performed with high-SNR surface coil quality, without patient re-positioning?
Max. scan range (without re-positioning of the patient, with high-SNR surface coil quality for the whole body):
Describe the technical solution for whole-body imaging:
13.1.2 | Whole-Body Applications
Whole-body MR angiography (with high-SNR surface coil quality, without patient re-positioning):
Whole-body metastases evaluation (with high-SNR surface coil quality, without patient re-positioning):
Can Parallel Imaging be applied throughout the whole body?
If not, state restrictions, i.e. for which body parts in a whole-body examination Parallel Imaging is not possible:
14 | Parallel Imaging
14.1 | Routine Parallel Imaging
A comprehensive package for Parallel Imaging shall be included.
An advanced package for Parallel Imaging shall be offered, with the functionality to enable '2-dim' Parallel Imaging, i.e. simultaneous PI in phase-encoding and 3D-encoding directions with 3D scans.
Simultaneous Parallel Imaging in 2 dimensions possible:
Is a high-speed Parallel Imaging technique available?
Is this Parallel Imaging technique compatible with the standard coil set?
Explicitly state possible restrictions, i.e. coils that are not compatible with Parallel Imaging:
Image-based algorithm for Parallel Imaging (e.g. SENSE):
k-space-based algorithm for Parallel Imaging (e.g. SMASH):
Is a separate measurement for coil sensitivity calibration necessary?
What is the time needed for the coil sensitivity calibration?
Is a new separate calibration measurement necessary if the patient and/or the patient table moves?
Does the software support the user in finding optimal parameter settings for Parallel Imaging?
Which parameters are taken into account by this online help function?
14.2 | Advanced '2-dim' Parallel Imaging
15 | STANDARD MULTI REGION Pulse Sequences . Each vendor to give his specific sequences.
15.1 | Spin Echo
15.2 | Fast Spin Echo
15.3 | Gradient Echo
15.4 | Inversion Recovery
15.5 | Phase Imaging
15.6 | Multislice/Multiecho
15.7 | 3D Volume Imaging
15.8 | Rectangular FOV
15.9 | Presaturation with auto-shim capability for fat suppression
15.1 | Minimum of 256 echo train length
15.11 | Fast Inversion Recovery
15.12 | Fluid Attenuated Inversion Recovery (FLAIR)
15.13 | Fast FLAIR
15.14 | Short Tau Inversion Recovery (STIR)
15.15 | Fast STIR
15.16 | Fast Gradient Echo
15.17 | MR Fluoroscopy
15.18 | 2DFT
15.19 | 3DFT
15.2 | FatSat
15.21 | Magnetization Transfer Contrast (MTC) and Ramped-RF (TONE)
15.22 | Motion Artifact Suppression Technique (MAST)
15.23 | No Wrap
15.24 | Phase Wrap
15.25 | Flow Compensation
15.26 | Gating to include peripheral pulse gating
15.27 | Echo Planar Imaging(EPI)
15.28 | Breathhold Cine,Fiesta (True FISP or Balanced FFE), for cine heart imaging, imaging of the veins
15.29 | ECG Triggering and ECG gating
15.3 | Imaging of the gastrointestinal tract, and is optimally suited for hi-res real- Time imaging
15.31 | Retrospective Gating
15.32 | T1 – FLAIR
15.33 | T1-weighted Phase Sensitive Inversion Recovery
15.34 | Double inversion recovery imaging
15.35 | balanced steady state free precession Spoiled Gradient Recalled
15.36 | Dual Echo Fast Gradient Echo
15.37 | IVIM and Kurtosis DWI
15.38 | Arterial Spin labeling technique
15.39 | Amide Proton transfer imaging (APT)
15.4 | 0 TE Options
15.41 | Dixon
15.42 | Spectral Water-Excitation, Fat saturation without increase of repetition time TR is Preferred
15.43 | TSE with Echo Sharing, dual contrast TSE (PD+T2), can be achieved without time penalty(Increase)
16 | Image Processor
16.1 | The image processor shall have hard disks dedicated to raw data ≥600GB; main memory ≥ 64 GB. ” the higher will be rated"
16.2 | Specify Image reconstruction speed in FFT/s or images /s.
17 | Computer System
17.1 | The computer system shall have a minimum of 64GB main memory (RAM) ” the higher will be rated"
17.2 | The computer system shall be of an open architecture windows design to allow for upgradability.
17.3 | System shall be able to communicate, send and receive images to a DICOM standard functions, PACS, HIS, RIS and/or teleradiology system or Any other applicable devices
17.4 | Enable remote service support
17.5 | Maximum Speed of image reconstruction, 2562 or 5122 without reducing FOV
17.6 | Color LCD flat screen monitor, 1024x 1024 matrix
17.7 | The array processor should be able to perform simultaneous reconstructions at the Same time during date acquisition.
18 | Image Storage
18.1 | Multiple Hard Disk drives of large capacity ≥ 2 TB for System Software and data storage. It shall have image storage capacity of at least 400,000 images or higher at 256x256 image matrix.
18.2 | CD-DVD RW for archiving and storage of images
18.3 | Compliant with all stand DICOM Specifications
18.4 | System should be cybersecure and patient data protected with the latest preventive patches.
19 | Accessories and other requirements
19.1 | MR compatible fire extinguisher
19.2 | Door mounted metal detector
19.3 | Non-Magnetic Wheel Chair:Wheelchair is constructed entirely of non-ferrous materials. Specifically designed for use in and around the MRI suite. All materials including fasteners and bearings are either non-ferrous metals or special polymers.
19.3 | Phantoms: For Image Quality audits and monitoring./ACR PHANTOM
19.4 | Phantoms: For Image Quality audits and monitoring,
19.5 | MR Compatible patient Music system
19.6 | CCTV SYSTEM WITH TWO WAY PATIENT with Camera Color and Zoom ( SCANNER ROOM) - OPERATOR ( CONTROL ROOM) INTERCOM SHALL BE INCLUDED
19.7 | Self-storing non-ferrous IV pole
19.8 | MR compatible Cabinet in the room for coils and accessories, Table pad, Positioning pads, Immobilization straps
19.9 | MR Compatible Automatic pressure Injector
20 | Automatic Injector: MR-compatible pressure injector for double barrel perfusion imaging, infusion pumps, contrast enhanced MR angiogram & bolus tracking. Name and model of injector offered to be mention clearly
21 | MRI Adjustable Height Stretcher
22 | Training
Onsite basic application training for all radiographers as needed
10 days Onsite intermediate application training for all radiographers
One week Onsite Advanced Cardiac,Post process and other special application training for senior radiographers and cardiac imaging radiologist.
One week to 10 day with interval period application training upon request
Service and maintenance training for biomedical engineers.
21 | Consumables
Should be Supplied with all standard consumables.
22 | Accessories
Should have All standard Accessories must be included as per manufacturer .
23 | Available Additional Option
No. | Options Description:
24 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and End user training for operators and Physicans is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Continuous Update and Upgrading of the system and all provided software with the system including and limited to the post processing and DICOM software, during the warranty (any hardware upgrade needed for such software upgrade should be included and provided with no additional cost by the vendor). It is expected from the vendor to always install the latest software versions within 4 months of its official release.
8 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
9 | All coils and MRI accessories, helium and Cold head, shillers should be full covered and included during warranty and maintenance contract periods.
10 | Maximum time needed for parts ordering and replacement (is calculated as a system down time )
11 | Maximum response time for system faults or breakdown calls (max. time for the engineer to be at the site) to be provided.
12 | Warranty must be minimum 5 years comprehensive from date of installation

## MRI Civil & MEP Works
Civil and MEP works for Open MRI ≥1.0 T
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Civil and MEP works for Open MRI ≥1.0 T | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
All Pre-installation requirements should be submitted, including suggested room Design and dimensions
1 | All room preparation and system installation requirements (to be clarified and agreed on after vendor initial site visit)- full room layout required.
2 | Scope of Proposed Design and As built Drawings ( after completion of Project)Should be Included
3 | Interior Desiging and Decorative works to be Praposed and should Include in scope of Works
4 | Demolition and Mobilization works should be included
5 | Mobalization(Wall opening and closing ) for MRI machine delivery should be included
6 | The system offered shall be designed to operate normally under the conditions of UAE, Environmental factors and MOH policy the conditions include Power Supply, Climate, Temperature, Humidity, Dust , etc
6 | RF cage:
6.1 | Supply and installation of RF cage includding RF doors and RF viewing window as per the praposed drawing. And full technical Specification of proposed RF cage
6.2 | RF cage Testing and Commssioning Certification should be provided
6.3 | RF Cage Integrity Testing for Compliance and report on all RF frequencies required
6.4 | Detailed test reports and compliance certification to be provided at each step to ensure the integrity of the RF shield not only meets but also exceeds the manufacturers requirements.
6.5 | Plan and include construction design must ensures an easily removed and re-fitting of wall/ceiling panels to allow magnet access during initial installation and in the event of an equipment upgrade.
6.6 | Provide overlap of the material at each panel joint for the maximum level of performance, attenuation integrity and electrical conductivity.
6.7 | To install appropriate internal fit out or lining, lighting and electrical meet 3.0T international standards.
6.8 | To meet the regulatory standards and safety of UHS campus the auxiliary components to be installed by the vendor, Items such as a Quench Pipe, Filter/Pen Panel, Assorted Waveguides, Medical Gases waveguides and Electrical Filters are mounted within the RF cage lining and positioned to suit site requirements. (vendor needs to share product certificate for these items)
7 | Water Chiller will be provided as per OEM recommendations
7.1 | Supplier should provide accurately the stable Scan Room Temperature and Humidity Control.
7.2 | To arrange variable air volume control system to monitor the temperature of the return air flow and adjusts the supply air flow to accurately maintain desired temperature within the scan room and humidity monitor
8 | MRI Safety/ Barrier/Warning/ RF Cage Doors (Zone 4)
8.1 | To provide trouble-free operation with no loss of attenuation capability as a result of friction and wear
8.2 | Magnetic door interlock
8.3 | Warning light and signs
8.4 | Cryogen venting(Quench) and emergency exhaust must be provided
9 | Partition works and Cealing works should be Similar to UHS standards and Compatable with open MRI . Proposed Specification need to be provided
9.1 | Aesthetic Ceiling.two quantity picture of ceiling and on wall picture can be done in MRI exam room or console room as per client suggestion
10 | Acoustic ceiling tiles.
10.1 | The maximum acoustic noise level expected outside the MRI room (state the technical specification of the door of MRI room offered - with its sound (acoustic noise) proof
10.2 | The Sound insulation can be incorporated into the framework of the walls during assembly and to ensure standard acoustic internal lining.
10.3 | Wall treatment
10.4 | To use Noise control and insulation materials only.
10.5 | To use only noise attenuating materials.
10.6 | To Fix the acoustic wall panels shall absorb 93% of the sound and report required.
10.7 | MR compatible multifunction lighting design and system.
11 | Quench Pipe’s warning board outside MR roomTo eliminate unrestrict and unsupervised access, the door to the scan room as below
12 | To provide clearest observation window with RF observation glass
12.1 | The observation Glass shall be anti-reflective coating
13 | Flooring- Medical Grade (similar or better than existing flooring at UHS)
13.1 | Removing of existing flooring & levelling with self-levelling screed for complete area.
13.2 | Proposal for Chiller Area Foundation including barriers should be provided ( Kindly have a site survey If current MRI chiller area is not Sufficent)
14 | Power Supply Capacity and Specifications
14.1 | Design of Power Layout
14.2 | Proposal for Supply and Instation of Electrical works. socket and switches should be via brand and of same type used in UHS premises.(All wires will be concealed and in conduct pipes and the wiring for electrical and network cable will be commensurate with the existing and MRI Compatable).
14.3 | The lighting should be all LED and should support energy conservation and open MRI Compatable
15 | HVAC- Study the current layout and Propose the modifications required
16 | Fire Alarm and Fire Fighting modfication works to be propose
17 | Supply and installation of MRI compatible Bedhead units(MRI room and patient receiving area )for proposed locations including Medical Gas Points and Nurse call system
Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
1
Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Equipment. Two Copies each and one Soft Copy for each.
3 | It is mandatory for the Contractor to submit End-of-Sale, End-Of-Support, and End-Of-Life for each individual hardware component - Proof documents from the manufacturer to be attached with the proposal.
4 | Proposal Should be submitted after having site survey and comparison with the current As built
5 | All The works Should Comply with Standards of Civil defence, SEWA and MOH
6 | Contractor should discuss the final technical proposal with the technical team before submission. Proposal should include ongoing warranty, support.