## Sheet1
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| MID LEVEL AUDIOMETER(DIAGNOSTIC) | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optionals . Please specify the details clearly and all the requested spec/info must be well supported with reference documents.
2 | The unit must be able to integrate with hospital HIS system . For that the softwares and hardware must be included and provided by vendor.
3 | Device must be with latest technology and standard of care
4 | Pure Tone :
4.1 | Must be having Air Conduction ; 125-8000Hz
4.2 | High Frequency ; 8000Hz to 20000Hz.( approx)
4.3 | Full Frequency Range ; 125Hz to 20000Hz( approx)
4.4 | Must be having bone Conduction ; 250Hz to 8000Hz ( approx)
4.5 | Sound field Range approx 125Hz to 12500Hz.
4.6 | Paired Inserts : 125Hz to 8000Hz.( Approx)
5 | HEARING LEVEL RANGE :
5.1 | For Air Conduction ; -10db HL to 120db HL
5.2 | High Frequency ; -10db HL to 100db HL
5.3 | For Bone Conduction :
5.4 | *-10db HL to 90db HL (Mastoid); -10db HL to 90db HL (Forehead)
5.5 | SIGNAL FORMAT must be having Steady , Pulsed , Frequency modulation (5Hz) and Pedeatric Noise.
5.6 | SPEECH audiometry must be available with all options.
5.7 | Machine must be provided with movable trolly with good standards.
5.8 | Masking option required with Narrowband Noise for both Air and bone conduction.
5.9 | Free Field testing should be available.
6 | SPECIAL TEST :
6.1 | ABLB , SISI , High Frequency Audiometry and Tone Decay to be included
6.2 | Quality and Patient Safety :
6.3 | Unit must be FDA approved/ CE Marked
6.4 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
6.5 | Unit must comply with Eletrical safety test IEC 60601-1 .
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for ENT/Audiology depr (End User Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .