## Indirect ophthalmoscope set
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Indirect ophthalmoscope set | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Wireless Headband Indirect ophthalmoscope
2 | Rechargeable battery operated
3 | Light weight, small & compact
4 | Should have internal metal chase for long life durability
5 | Adjustable hand band circumference from: 530 to 630 mm
6 | Able to view the disc, macula and eye vessels
7 | LED Adjustable illumination
8 | Should allow stereo views in pupils ranging from 1mm to 10mm.
9 | Mirror adjustment for angle and height
10 | Overhead locking knobs for optical unit adjustment
11 | Single control for changing beam site and position
12 | Wide angle diffuse
13 | Rubber eye cap
14 | Eyepiece: 1.5 diopter
15 | Diffuser filters to expand the illumination beam and allows enhanced viewing peripheral areas of the retina
16 | Full range of Apertures (Large Aperture - Intermediate -Small Aperture)
17 | Full range of filters: yellow, cobalt blue, red-free and clear
18 | UV and IR protection
19 | Wall mounted charging unit with extra spare battery
20 | Should come with 20 D lens
11 | Device must be manufactured on good standards like Qulaity management system , ISO-13485
12 | Offered machine must be having good safety features , equipped with safety alrams , technical alrams and system interlocks.
13 | Offered unit -History of incident,recalls and safety alerts .
14 | FDA approved / CE Marked
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .

## Retinoscope Set
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Retinoscope Set | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | For a fast, reliable and objective determination of refractive error
2 | Polarization filter: Eliminates stray light and internal reflexes
3 | Dustproof housing
4 | Brilliant, precise streak image
5 | ParaStop function for the precise selection of the parallel beam
6 | Enclosed vergence sleeve prevents drift during examinations
7 | LED illumination
8 | Patented continuous brightness control
9 | Fixation cards with holder for dynamic retinoscopy
10 | ParaStop features the latest multi-coated optics for an exceptionally bright fundus reflex and easier
11 | detection of the neutralization point
12 | Device must be manufactured on good standards like Qulaity management system , ISO-13485
13 | Offered machine must be having good safety features , equipped with safety alrams , technical alrams and system interlocks.
14 | Offered unit -History of incident,recalls and safety alerts .
15 | FDA approved / CE Marked
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .

## 3in1 AutoRef
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| All in one Autoref kerato , tonometer , Pachymetry | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
System must be having Egonomic design and easy to use
A | Space saving autoref , kerato , noncontact tonometer and pachymeter
B | Large pupil zone imaging method
C | Super luminescent diode (SLD) light and highly sensitive CCD camera
D | Optimal fogging to minimize accommodation
E | Double mire ring method
F | Automated calculation of corrected IOP
G | Easy to use 7 inch screen
H | 3D auto tracking and auto shot
I | Should have joystick for flexible alignment
1 | Auto refractometer
Measurement range :
A | Auto keratometer
B | Measurement range
C | Curvature radius 5.00 to 13.00 mm (0.01 mm increments)
D | Refractive power 25.96 to 67.50 D (0.12 / 0.25 D increments)
E | Cylindrical power 0 to 12.00 D (0.12 / 0.25 D increments)
F | Axis 0 to 180º (1° / 5° increments)
G | Pupil size measurement range (Angle ,2.0mm)
H | 1.0 to 10.0 mm (0.1 mm increments)
I | PD measurement range ( 1mm , 20 to 85mm)
J | Must be having auto fog system
2 | NON-CONTACT TONOMETER :
A | Measurement range
B | 1 to 60 mmHg (1 mmHg increments)
C | Working distance
3 | NON CONTACT PACHYMETER
A | Measurement range
B | 300 to 800 µm (1 µm increments)
4 | Printer
A | Preferably : Thermal line printer with easy loading and auto cutter
5 | Chinrest travel distance :
A | Up/down : 65- 70mm
6 | System should not bulky , Please specify the weight and unit must be provided with good quality trolly .
7 | System must be having Electrical safety test IEC 6060-1
8 | Device must be manufactured on good standards like Qulaity management system , ISO-13485
9 | Offered machine must be having good safety features , equipped with safety alrams , technical alrams and system interlocks.
10 | Offered unit -History of incident,recalls and safety alerts .
11 | FDA approved / CE Marked
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .

## CompLensmeter
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| COMPUTERIZED LENSMETER | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Portable , light weight and user friendly design
2 | Preferably Hartmann sensor with more then 120 measurement points ; which ensures easy measurements of all types of lenses ; including progressive and contact lenses.
3 | Automatic lens type detection
4 | Prism layout function
5 | Full graphic vertical minimum 5.7-inch color LCD with touch screen
6 | UV transmittance measurement
7 | Refractive measurement
8 | High-speed line printer with auto cutter
9 | Measurement range :
A. | Sphere (Spectacle lenses): -25.00 to +25.00 D
B. | Sphere (Contact lenses): -25.00 to +25.00 D (BC=6.0 to 9.0)
C. | Cylinder: 0.00 to ±10.00 D (-, MIX, +)
D. | Axis: 0 to 180º (1º increments)
E. | Prism: 0.00 to 20.00Δ
F. | Prism mode: Δ, Base In / Out, Base Up / Down
G. | PD measurement : 15 to 45 mm approx (monocular), Single vision PD,Progressive lens far vision PD
H. | UV transmittance: 0 to 100% (minimum 1 or 5% increments) with central wavelength around 365 nm (UV-A)
I | Measurable lens diameter
J | Spectacle lenses preferably - ø20 to 120 mm
K | Contact lenses Larger than the inner diameter of the nosepiece (ø5 mm)
L | Compensation function for high index lenses: The Abbe number is changeable in the range of 20 to 65.
M | Marking system: Ink cartridge type/Thermal .
10 | System must be having Electrical safety test IEC 6060-1
11 | Device must be manufactured on good standards like Qulaity management system , ISO-13485
12 | Offered machine must be having good safety features , equipped with safety alrams , technical alrams and system interlocks.
13 | Offered unit -History of incident,recalls and safety alerts .
14 | FDA approved / CE Marked
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .

## LED Chart
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| LED CHART SYSTEM | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Testing distance
2 | Viewing area W*D(mm)
3 | Visual Acuity range/Steps
4 | Standrad test ; Please specify with the list
5 | Binocular test ; Please specify with the list
6 | Unit should be light weight based on latest technology and easy to use .
7 | Device must be manufactured on good standards like Qulaity management system , ISO-13485
8 | Offered machine must be having good safety features , equipped with safety alrams , technical alrams and system interlocks.
9 | Offered unit -History of incident,recalls and safety alerts .
10 | FDA approved / CE Marked
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .

## Opthalmic Unit
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| OPTHALMIC UNIT WITH AUTOMATED PHOROPTER AND FULL APERTURE TRIAL LENS SET | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | OPTHALMIC UNIT :
A | Sliding Two instruments tabletop
B | Should have motorized up and down tabletop (708 mm to 908 mm)
C | Should have 3 drawers cabinet
D | Should have control panel
E | Should have column with adjustable LED lamp
F | Should come with reclinable rotatable chair (60° reclining, 45° left 45° right rotation)
G | Should have reclinable armrest and footrest
H | Should have wheelchair access capability by movable base
I | Should have padded headrest to keep patient comfortable
2 | Automated phoropter
A | Should have min 5.7-inch Color LCD control screen
B | Advanced Dial with S/C/A Mode Keys
C | Should have Cross Cylinder Examination
D | Should have Built-In High-Speed Line Printer with Automatic Paper Cutting
E | Ergonomically Designed Refractor Head for Comfortable Examinations
F | High Quality Near Point Card
G | Easy-to-Clean: Detachable Forehead Rest, Face Shields, and Protective Glasses with Antifouling Coating
Measurement range
A. | Sphere: Preferably (-19.00 to +16.75 D) (0.25 / 0.50 to 3.00 D increments)
B. | Cylinder: Preferably - 0.00 to ±6.00 D (0.25 / 1.00 D increments)
C. | Axis: Preferably - 0 to 180° (1° / 5° increments)
D. | PD: 48.0 to 80.0 mm (far mode), 50.0 to 74.0 mm (near working distance of 35 cm), 54.0 to 80.0 mm (far PD possible for 100% convergence)
E. | Rotary prism: 0.0 to 20.0Δ (0.1 / 0.5 / 2.0Δ increments)
Auxiliary lenses
A. | Cross cylinder lens: ±0.25 D
B. | Occluder (included)
C. | Pinhole plate (ø2.0 mm)
D. | Red maddox rod (Right eye: horizontal, Left eye: vertical)
E. | Red / Green filter (Right eye: red, Left eye: green)
F. | Polarizing filters (Right eye: 135°, Left eye: 45° / Right eye : 45°, Left eye : 135°)
G. | Fixed cross cylinder lens (±0.50 D)
3 | Full Aperture Trial Lens Set in vinyl case (250 + lenses)
Trial Lens Sets in cover that should fit in opthalmic machine drawer , Wide range of lenses must be available that must be easy to fit and remove from trial frames with clear markings and a removable tray. The trail lens outer cover (frame) must be mettalic rather then plastic .
4 | System must be having Electrical safety test IEC 6060-1
5 | Device must be manufactured on good standards like Qulaity management system , ISO-13485
6 | Offered machine must be having good safety features , equipped with safety alrams , technical alrams and system interlocks.
7 | Offered unit -History of incident,recalls and safety alerts .
8 | FDA approved / CE Marked
3 | Accessories
Supplied with all standard accessories.
4 | Consumables
Supplied with all standard consumables.
5 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
6 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .

## Portable Slit Lamp
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Portable Slit Lamp | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Should be LED powered
2 | Should have at least 20,000 LUX LED power source
3 | Should be powered by AAA Ba!eries ..
4 | Magnification - 10X and 16X
5 | Objective lens 80mm
6 | Working distance ; 100- 110mm
7 | Field of view 10mm and 15 mm
8 | Ocular visibility -8 to + 5
9 | Slit width .01 x 12 mm to .08 x 12mm
10 | Spot diameter 1mm to 12mm
11 | Built in blue filter
12 | Device must be manufactured on good standards like Qulaity management system , ISO-13485
13 | Offered machine must be having good safety features , equipped with safety alrams , technical alrams and system interlocks.
14 | Offered unit -History of incident,recalls and safety alerts .
15 | FDA approved / CE Marked
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .

## Slit Lamp
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| SLIT LAMP | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | SLIT LAMP :
A | Slit-lamp to perform routine eye examination for diagnostic of ocular pathologies, injureis and for posterior segment examination
B | Heavy slitlamp base for better support and accurate movement
C | Microscope drum should be binocular with 5 magnifications (6×10×,16×,25×,40×)
D | Desired Slit beam width (0-14 mm) and slit beam length (1-14 mm) with aperture diameter starting from 0.2mm - 14
E | Should have high light transmission and the optical excellence ensure a superb clinical view
F | Should have LED light source provides even illumination with the sharpest and most homogenous slit in the market eliminating any filament 30.000 Hours life time
G | Should be capable of rotating 0 to 180 degree and adjustable brightness
H | Should have the integrated filters (Amber filter , grey, red free , blue)
I | Should have cross slide mechanism Perfection in design, materials and construction allow these movements to be made smoothly and effortlessly without any wrist or finger strain – and to continue doing so for the very long lifetime of this classic slit lamp.
J | Should have applanation tonometer for the measuring of intraocular pressure with the option for disposable prisms
K | Should come with integrated imaging system with integrated beam splitter and high-resolution digital camera along with the operation software and all in one pc
L | Imaging system should have control panel Located in front of the joystick, the control panel can be operated blindly while examining the patient. It provides not only full control over the slit lamp illumination, but also a release bar for the camera, as well as simple and ergonomic management of the camera settings.
M | Camera should be minimum 5 um pixel size for clear and crisp image
N | Camera is equipped with aperture control on five levels for more control on the image needed
O | Camera should have high frame rate per second
P | Should be dicom compatible with no extra cost
Q | Automated detection of right and left eyes
R | Chinrest Unit : Amount of Vertical movement of chinrest around 75-85 mm
S | Unit must be provided with good base table, ( Please specify the details in your offer)
2 | System must be having Electrical safety test IEC 6060-1
3 | Device must be manufactured on good standards like Qulaity management system , ISO-13485
4 | Offered machine must be having good safety features , equipped with safety alrams , technical alrams and system interlocks.
5 | Offered unit -History of incident,recalls and safety alerts .
6 | FDA approved / CE Marked
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors , nursing staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .

## FundusImaging
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| ULTRA WIDE FIELD FUNDUS IMAGING SYSTEM | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Top of the range, Compact ultra-wide field retinal imaging system, which supports the diagnosis, analysis, monitoring and documentation of ocular pathology in the periphery
2 | Should be able to produce a 200° captures more than 80% of the retina in a single image
3 | Fast exposure time: ≤ 0.5 seconds
4 | 3-in-1 color Depth imaging provides important clinical data from the retinal surface through the choroid.
5 | Should be cSlO technology images through most cataracts and
6 | Imaging modes: Fluorescein angiography (fa), auto-fluorescence (af), red free, Choroidal and Color
7 | Should have image overlay tool facilitates comparison of images in different image modes and from visit to visit.
8 | Should have Auto montage combining series of pictures into a single image showing up to 220° (97%) of the retina.
9 | Comprehensive retinal analysis software package, with advanced image review platform
10 | Should have arm rest for FA procedures to insure patient comfort
11 | Latest PC version with minimum 17” HD LED Monitor, 4 TB HD, 8GB RAM, OS Windows 10 or higher, DVD-R/RW , USB interface
12 | Desktop laser jet color printer
13 | Mounted on motorized table
14 | System must be having Electrical safety test IEC 6060-1
15 | Device must be manufactured on good standards like medical device Qulaity management system , ISO-13485
16 | Offered machine must be having good safety features , equipped with safety alrams , technical alrams and system interlocks.
17 | Offered unit -History of incident,recalls and safety alerts .
18 | FDA approved / CE Marked
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors, nursing staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .