## Sheet1
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| All in one Autokerato Refractometer, tonometer and pachymetry | 2 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
| System must state of the art , latest release model , having Egonomic design and easy to use |  | Yes
A | Space saving autoref , kerato , noncontact tonometer and pachymeter
B | Large pupil zone imaging method
C | Optimal fogging to minimize accommodation
D | Automated calculation of corrected IOP
1 | Auto refractometer
Measurement range :
1.1 | Auto keratometer ( Please specify the in built technology)
1.2 | Please specify the Measurement range
1.3 | Curvature radius 5.00 to 13.00 mm (0.01 mm increments)
1.4 | Refractive power 25.96 to 67.50 D (0.12 / 0.25 D increments)
1.5 | Cylindrical power 0 to 12.00 D (0.12 / 0.25 D increments)
1.6 | Axis 0 to 180º (1° / 5° increments)
1.7 | Pupil size measurement range (Angle ,2.0mm)
1.8 | 1.0 to 10.0 mm (0.1 mm increments)
1.9 | PD measurement range ( 1mm , 20 to 85mm)
1.91 | Must be having auto fog system
2 | NON-CONTACT TONOMETER :
2.1 | Measurement range
2.2 | 1 to 60 mmHg (1 mmHg increments)
2.3 | Please specify the Working distance
3 | NON CONTACT PACHYMETER
3.1 | Please specify the Measurement range
3.2 | 400 to 750 µm (1 µm increments)
4 | Printer
4.1 | Preferably : Thermal line printer with easy loading and auto cutter
5 | Chinrest travel distance :
5.1 | Up/down : 65- 70mm
6 | Quality and Patient Safety :
6.1 | FDA approved / CE Marked
6.2 | System should not be bulky , Please specify the weight and unit must be provided with good quality trolly .
6.3 | offered System must comply with Electrical safety IEC 6060-1
6.4 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .