## MOTOMED
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Legs and arms Trainer | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the body trainer and optional items .Please specify the details clearly and all the requested specification info must be well supported with reference documents.
2 | Please specify make, model and origin of the machine
3 | Touch screen display ,retractable and inclinable
4 | Soft Grip handles
5 | 2-stage pedal radius adjustment (7 cm to 12.5 cm)
6 | Pedal axle of the arm trainer is height adjustable from 90 cm to 105 cm
7 | Pedal axle of the leg trainer is height adjustable from 30 cm to 45 cm
8 | Plastic-coated safety foot shells with Velcro fastening
9 | Easily movable with large transport rollers
10 | Therapy and motivation programs(Active/Passive/assited/resistance modes)
11 | Closed Metal/plastic body (PC/ABS), easy to clean and disinfect.
12 | Maximum permissible user weight
13 | Specify Dimensions in cm
14 | specify Weight of the unit
15 | Must have universal power adapter 240v ,50/60HZ
16 | Quality and Patient safety
Unit must be FDA / CE marked .
Unit must comply ISO 13485 Quality management system for medical device or equivalent .
System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1
17 | Accessories
Supplied with all standard accessories.
18 | Consumables
3.1 | Supplied with all standard consumables.
19 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
20 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and End user training for nursing staff.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
E | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
F | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
G | Warrenty : Comprehensive minimum 3 Years warrenty .

## COMBINATION UNIT
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Combination Therapy unit (Electrotherapy and Ultrasound) | 2 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the body trainer and optional items .Please specify the details clearly and all the requested specification info must be well supported with reference documents.
2 | Please specify make, model and origin of the machine
3 | Combination therapy (treating one injury simultaneously using a combination of electro and ultrasound therapy
4 | Full color touch screen navigation
5 | Color guided treatment method visualization
6 | Should have Electro therapy, 2 and 4 poles, 4 channels completely independent
7 | Interferential currents( Isoplane/dipole vector field/Classical Interferiantail)
8 | Tens electrotherapy/Diagnostic programs/Simultaneous /stimulation therapy
9 | Must have Pencile electrode with sponge
10 | S-D curve diagnostic programs
11 | Micro current (Modulated and surge)
12 | High Voltage (HVPC)(surge)
13 | 2 ultrasound output connectors
14 | Must have Multifrequency ultrasound head (1 and 3 MHz), 5 and 1 cm²
15 | Continuous and pulsed mode (10–20–30–40–50-100%)
16 | Acoustic and visual contact control led
17 | Must be having two section trolley
18 | Specify Dimensions (H x B x D)cm
19 | specify Weight of the unit
20 | Must have universal power adapter 240v ,50/60HZ
21 | Quality and Patient safety
A. | Unit must be FDA / CE certified
B. | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
C. | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1
22 | Accessories
Supplied with all standard accessories.
23 | Consumables
Supplied with all standard consumables.
24 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
25 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and End user training for nursing staff.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
E | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
F | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
G | Warrenty : Comprehensive minimum 3 Years warrenty .

## chiller unit (Coldpack)
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Physiotherapy cold Pac unit (Chilling unit) | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the cold pack unit and optional items .Please specify the details clearly and all the requested specification info must be well supported with reference documents.
2 | Please specify make, model and origin of the machine
3 | Cold therapy applied directly to the affected area of the patient and control pain due to inflammation, swelling, multiple sclerosis, muscle strains, sprains and spasms.
4 | Effective relief of pain
5 | The temperature of the cold pack up to -12 °c
6 | Must include FlexiPacs small ,Medium and large with latex free
7 | unit must have temperature monitor
8 | Closed-cell foam insulation for efficiency
9 | Stainless steel body
9 | Heavy-duty compressor for more efficient cooling
10 | Specially designed interior chilling coils cool packs faster
11 | 8 cm (3”) swivel-type rubber casters for silent, friction-free movement of mobile units
12 | Drain valve for easy cleaning and defrosting
13 | Ny latex wraps for cold packs
16 | Unit Must be available for demonstration
18 | Specify Dimensions (H x B x D)cm
19 | specify Weight of the unit
20 | Must have universal power adapter 240v ,50/60HZ
21 | Quality and Patient safety
A. | Unit must be FDA / CE certified
B. | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
C. | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1
22 | Accessories
Supplied with all standard accessories.
23 | Consumables
Supplied with all standard consumables.
24 | Available Additional Option
No. | OPTIONS DESCRIPTION
25 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and End user training for nursing staff.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
E | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
F | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
G | Warranty : Comprehensive minimum 3 Years warranty .

## Bobath couch
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Bobath treatment couch | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the Bobath unit and optional items .Please specify the details clearly and all the requested specification info must be well supported with reference documents.
2 | Please specify make, model and origin of the machine
3 | Bobath treatment techniques require a suitably wide treatment surface with a safe and capable lifting capacity for both patient and therapist.
4 | Couch must have a safe working load with three section
5 | Operation of the bed must be electric
6 | Patient weight load up to 250 to 300 kg
7 | Foot paddle switching functions
8 | Paper Roll Holder
9 | heavy duty Castor Base
10 | Specify total Length of the unit in ___cm
11 | Specify Width in___cm
12 | Specify Backrest length in __cm
13 | Specify Long section length in ___cm
14 | Adjustable backrest ___ Degrees
15 | Height range from 49cm to 93.5cm in approx. 30 seconds
16 | Foam Density ___kg/m3
17 | Specify Dimensions (H x B x D)cm
18 | specify Weight of the unit
19 | Must have universal power adapter 240v ,50/60HZ
20 | Quality and Patient safety
A. | Unit must be FDA / CE certified
B. | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
C. | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1
21 | Accessories
Supplied with all standard accessories.
22 | Consumables
Supplied with all standard consumables.
23 | Available Additional Option
No. | OPTIONS DESCRIPTION
25 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and End user training for nursing staff.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
E | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
F | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
G | Warranty : Comprehensive minimum 3 Years warranty .

## Gait trainer with hoist
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
Gait Trainer with hoist for Inpatient |  | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the Gait trainer and optional items .Please specify the details clearly and all the requested specification info must be well supported with reference documents.
2 | Please specify make, model and origin of the machine
3 | GaintTrainer enables partial weight -bearing therapy to be conducted with the assurance of patient Comfort and Safety, and with convenient access to the patient for manual observation and assistance
4 | Electrically motarized, the unit must be with mains operated built in patient hosit .
5 | Single point, with 360 degree rotational bar to help patient turn easily.
6 | Weight Capacity up to 150 -250 kg
7 | Hoist slings- Small,Medium and large.
8 | Easily disassembles for transport and storage
9 | Critical measurements-Floor to Hip, Floor to Chest, Trunk Width.
10 | Closed Metal/plastic body (PC/ABS), easy to clean and disinfect.
11 | Maximum permissible user weight
12 | Specify Dimensions in cm
13 | specify Weight of the unit
14 | Must have universal power adapter 240v ,50/60HZ
15 | Quality and Patient safety
A. | Unit must be FDA / CE marked .
B. | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
C. | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1
16 | Accessories
Supplied with all standard accessories.
17 | Consumables
Supplied with all standard consumables.
18 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
19 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and End user training for nursing staff.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
E | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
F | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
G | Warranty : Comprehensive minimum 3 Years warrenty .

## Rehabilitation Robot Gloves
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Rehabilitation Robot Gloves | 1 Pair | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the Robotic Gloves and optional items .Please specify the details clearly and all the requested specification info must be well supported with reference documents.
2 | Please specify make, model and origin of the unit
3 | Finger Rehab Training Device, Rehabilitation Robot Gloves left and right with Touch Screen, Finger Trainer Finger Extension Exercise for Long-Term Relief, for Stroke Hemiplegia Patients with Hand Dysfunction
4 | Help patients relieve hand spasms and stiffness by training as follows:
4.1 | Passive Training
4.2 | Pinch Training
4.3 | Finger Mirror Training
4.4 | Extend Training
4.5 | Functional Training
4.6 | Fine Motor Training
5 | Color guided treatment method visualization
6 | The glove can buffer and protect the metacarpophalangeal joint and the interprocedural joint
7 | Specify Dimensions (H x B x D)cm
8 | specify Weight of the unit
9 | Must have universal power adapter 240v ,50/60HZ
10 | Quality and Patient safety
A. | Unit must be FDA / CE certified
B. | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
C. | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1
11 | Accessories
Supplied with all standard accessories.
12 | Consumables
Supplied with all standard consumables.
13 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
25 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and End user training for nursing staff.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
E | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
F | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
G | Warrenty : Comprehensive minimum 3 Years warrenty .