## Vascular Doppler
Indicative Specifications / Evaluation Form
Item No. | POCT Ultrasound for OT/Anesthesia dept | Qty. | To Be Filled By Vendor | To Be Filled By Hospital
|  | 1 | Supporting Documents Reference | Met
| Technical Specs |  |  | Yes | No
Portable state of the art POCT Ultrasound Machine for the operating theater/ Anesthesia care.
Frame/design
1 | The system should fulfills the requirement of an easy-to-use, space-saving, and portable imaging system. and to be compact design for a better fit in the existing clinical environment of OT .
2 | Image acquisition from POCUS utilization by anesthesiologists involves the assessment of multiple organs in different perioperative situations.
3 | Should has User-friendly interface, preferably Touchscreen, Capacitive Technology for easy and quick settings adjustment
4 | Imaging modes : B , B+B: Dual , B+ Anatomic M, CFM: Color Doppler , PW: Pulsed Wave, Doppler / Power Doppler
5 | With atleast three Probes/Transducers for POCT/OT/Anesthesia use , including two linear array probe including high frequencyand one convex probe.
6 | Needle guide assessment for POCUS ultrasound
7 | The System should be on a compact Height Adjustable Trolley with Lockable Castors and Probes/Transducers hanger.
8 | Preferably more or less 15" Touchscreen, Capacitive, with Anti-glare and Wide Viewing Angle
9 | Built in Rechargeable Battery for minimum back-up use of 90-120 min.
10 | DICOM and HL7 with Wireless communication capability.
11 | Storage for Still Image and Loops.
12 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optionals . Please specify the details clearly and all the requested spec/info must be well supported with reference documents.
13 | Unit must be able to integrate with hospital HIS system . For that the softwares and hardwares must be provided by vendor .
14 | Unit must be FDA approved/ CE Marked
15 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
16 | Unit must comply with Eletrical safety test IEC 60601-1 .
2 | Accessories
2.1 | Vendor to quote supply all mandatory accessories that will required for full intended functioning of the equipment
3 | Consumables
3.1 | Vendor must include prices of all standard and optional consumables.
4 | Available Additional Options
No. | Options Description
Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for Nurses and Physicians (End User Training) is mandatory.
3 | 60601-1-1 for Medical Electrical Equipment - Part 1: General Requirements for Safety 1: Collateral Standard: Safety Requirements for Medical Electrical Systems
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Should have local service facility .The service provider should have the necessary equipment recommended by the manufacturer to carry out preventive maintenance test as per guidelines provided in the service/maintenance manual
7 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
8 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
9 | Warranty: 3 years comprehensive