## Interventional radiology system
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Interventional radiology (IR)system | 1 | Comply Yes/No (If Yes Specify) | Supporting Documents Ref. | by Hospital
Specifications
The start -of -the-Art Digital Interventional radiology (IR) system, Zero Gravity radiation Protection system,advanced high-end latest model.Biplane imaging allows doctors to follow the path of blood flow through the Patient's vessels and to create a "roadmap" for reaching and treating the precise location of disease or malformation. Floor & Ceiling mounted system, with better maneuverability & flexibility in positioning. Facility for motorized & manual positioning / rotation of stand from the floor pivot by +/-135 degrees for improved workflow and for ease of operation from both left and right side of the patient in addition to zero-degree normal head end position for floor mounted Gantry. Patient access should be possible from either left or right side with full patient coverage. Gantry should move at approx.18 deg/sec for angulation & approx. 25 deg/ sec for rotation with non-contact collision sensing collision protection mechanism working on capacitive sensing technology that works on actual body size to protect the patient from unexpected contact between the detector and the body.
2 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
3 | All pre-installation requirements should be submitted along with the quotation including suggested room dimensions etc.
4 | Integration with PACS , GE Dose Watch & RIS :Unit must be able to integrate with hospital PACS system & GE Dose Watch .All softwares and hardwares must be provided by vendor to do the intergation .
5 | DICOM Image Storage Service Class as defined below: The DICOM Storage Commitment Service Class For retrieving prior images from the PACS, the DICOM Query/Retrieve (C-FIND and C-MOVE for Patient Root and/or Study Root Model) Service Class should be support.
6 | Post-anterior Gantry & Lateral Gantry:
6.1 | Ceiling mounted Fully Motorized C-arm for the Post-anterior Gantry
6.2 | Floor/Ceiling mounted Fully Motorized C-arm for the Lateral Gantry
6.3 | Projection angles: 45º cranial to 45º caudal rotation
6.4 | C-Arm Rotation 120º LAO to 120º RAO
6.5 | Source-Detector distance (90-120 cm)
6.6 | Depth 90 cm or more
6.7 | Automatic position controller
6.8 | oblique angles for anterior versus posterior relationship Check / LAO/RAO projections
6.9 | In Bi-plane mode, straight caudal and straight cranial setting should be possible.
6.1 | Flex Arm with Flexibility in C-Arm movement for all kinds of procedures
7 | Table Top:
7.1 | Motorized 4 Way tabletop with comfortable mattress
7.2 | Floor-mounted swiveling patient table with telescopic foot and floating tabletop.
7.3 | Tabletop (narr.) /Mattress(thin)
7.4 | Carbon fiber patient positioning tabletop narrow including special-foam mattress. Mattress including cover.
7.5 | Foot Switch Monopl.(Wireless)
7.6 | For release of fluoroscopy, exposure and table brake as well as a configurable additional function. Wireless connection via radio communication.
7.7 | Table controls attached to the table side
7.8 | Table height (min.-max.) approx. 74 cm - 102 cm
7.9 | Tabletop length approx.___ cm
7.1 | Tabletop width approx.___ cm
7.11 | Longitudinal float range approx. ___ cm
7.12 | Lateral float range approx. 36 cm
7.13 | Table up/down speed approx. 30mm/s
7.14 | Separate Free table top dedicated use for IR radiologist (Top table on the examination table top)
7.15 | Patient weight 280 kg or more
8 | X-Ray Generator & Tubes and collimators
8.1 | High frequency, three phase microprocessor controlled with power output of 100 KW
8.2 | kVp range: 50 to 125 kV or equivalent .
8.3 | mA range: 3 to 1250mA or equivalent
8.4 | mAs range: 0.5 to 600mAs or equivalent
8.5 | Different levels of dose selection should be possible with dose selection at table side.
8.6 | Automatic Exposure Control
8.7 | Dose Area Product Meter (DAP)
8.8 | Dose Reduction Capability
8.9 | Latest technology of rotating anode x-ray tube
8.1 | Latest technology Dual focus grid pulsed X-ray tube with high anode heat storage capacity (>2 MHU), heat dissipation rate (> 10 KW/min) and stable focal spot or equivalent values.
8.11 | Semitransparent filters
8.12 | Possibility for independent rotation & off centre positioning
8.13 | Possibility of collimator positioning on last Image hold
8.14 | Automatic iris collimation
8.15 | Beam hardening filters of 0.2mm, 0.5mm and 1.0mm or equivalent to reduce stray radiation
9 | Digital Flat Panel detectors
9.1 | Two High resolution Flat detectors
9.2 | spatial resolution 3.0lp/mm or better
9.3 | Detector quantum efficiency of > 75% or better
9.4 | Integrated collision protection mechanism
9.5 | Patient acquired dose monitoring and display
9.6 | Handle for easy holding it
9.7 | Digital acquisition technology and digital subtraction angiography in matrix 1k.CLEARstent enables an improved display of vascular supports (stents).
9.8 | Angle-triggered native rotation angiography.
9.9 | plane-A 30x40 cm Dynamic Flat Detector subsytem for fluroscopy and cine-fluorography or equivalent
9.1 | 30x40 cm for good coverage, it must reach as close as possible to the anatomy(head) to avoid scattered radiation, reduce magnification & reduce dose, with approx. 1560 x 1440 pixels at 16-bit depth, with approx. 7 formats of zoom. Source Image Distance range: 87.4 cm -130.2 cm. Resolution: at least 2.7 lp/mm
9.11 | Plane-B 40x40 cm triple mode Dynamic Flat Detector or equivalent
9.12 | The outer detector physical housing is __cm diagonal square
9.13 | The digital output of the Flat detector is a 1024x1024 matrix at 14 bit depth.
9.14 | The pixel pitch is <160 micron, lower pixel pitch preferred for better resolution.
10 | Fluoroscopy system:
10.1 | Programmable modes of fluoroscopy & acquisition
10.2 | Pulsed fluoroscopy with grid control
10.3 | Pulse width should be manually and automatically adjustable
10.4 | Pulse rate of ≥30 PPS
Ability to provide a perspective roadmap for drawing any annotation
10.5 | Anatomic programming
10.6 | Console to display post-exposure technique values (kVp, mAs) for fluoroscopic and digital cine acquisitions
10.7 | fluoroscopic guided access for different oblique views
10.8 | Air gap technique, higher kVp, pulsed fluoroscopy, and large image receptor with collimation
10.9 | Automatic brightness control and Must have fluoroscopy road map function that automatically extracts blood vessel margins from DSA images.
10.1 | Wireless foot switch to control fluoroscopy and acquisition.
11 | Digital System & Image Pre Processing control Console:
11.1 | Acquisition settings, Automatic position control, Image processing, quantitative analysis, direct control of exam and run cycle, contrast, brightness and edge enhancement , hemodynamics system control and Valve sizing should be possible from the console [touch panel] at the tableside.
11.2 | Different frame Cardiac acquisition rates
11.3 | Acquisition, storage and archival with the available offered matrix
11.4 | Include all the advanced software applications for measurements such as calculations pertaining to Stenosis, diameter of catheter, Valve sizing etc. applicable for various kinds of patients.
11.5 | Facility to integrate IVUS,OCT, & FFR images in the lab
11.6 | C-arm control module and collimator control module shall be able to be installed separately.
11.7 | Equipped with the function that C-arm swivel from head position when the table top is moved to head direction and C-arm revert to head position when the table top is moved to foot direction
11.8 | C-arm inner diameter shall be 945mm or more.
11.9 | High storage Hard disk capability
12 | Display Monitors with Suspension:
12.1 | Monitor suspension with motorized height adjustment with possibility to move the monitor suspension from left side to the right side.
12.2 | Image system interface to the displays in the control room if the image system is installed in the control room.
12.3 | Control room (3 no’s): 1 no 21” or higher HR B/W flat screen monitor with high contrast ratio for images & 2 no’s 21” or higher HR Flat panel monitors Color monitors 1 each for display of haemo dynamic data / 3D CA data and another for patient demographics etc.
12.4 | Examination Room: Single large screen with all applications displayed on it. .Display suspension system flat displays with ≥50" monochrome flat displays with good background color.
12.5 | Roadmap functionality and dual reference function.
12.6 | Uninterruptible power supply with battery backup for 220-230 V
12.7 | CD/ DVD writing and printer facility and DICOM software for review of images. The DICOM media [CD/DVD] written using this system should be readable on personal computer with necessary DICOM software coming up automatically
12.8 | Standby power UPS 220-230V
12.9 | Uninterruptible power supply with battery backup for 220-230 V
13 | Master SW Package
13.1 | System SW Master package assists the Interventional radilogy doctor , technologists and catheterization lab managers with comprehensive data access and workflow-oriented tools in their daily reporting and administrative tasks.
13.2 | Automap
13.3 | Automatic stand positioning depending on the selected reference image and automatic reference image selection depending on the stand positioning.
13.4 | Data such as the dose-area product can be transferred to the RIS(GE dose Watch ).
15 | LB rad. protection w/ pivot arm
15.1 | For shielding the lower body against scattered radiation within the examiner’s moving range.
15.2 | Specially designed for avoiding collisions with the tube during oblique projections, therefore especially suited for cardiology.
15.3 | Upper Body Rad. Prot.
15.4 | To protect the upper body against scattered radiation within the operating range of the examiner, e.g. during interventional procedures.
16 | Examination lamp, 230 V
16.1 | Ceiling-mounted, flexible OR lamp that is adjustable toward the operator, to be used for diagnosis and minor surgery.
16.2 | Examination lamp with focusable dielectric light system.
16.3 | Light intensity: 50,000 lux (3,255 fc) at a distance of 100 cm
16.4 | Working distance: 70 to 140 cm
16.5 | Color rendering index Ra (gen.): 96
16.6 | Color temperature: 4,300 Kelvin
16.7 | Focusable light field size: 14 to 25 cm
16.8 | Lamphead diameter: 22 cm
16.9 | Halogen bulb: 22.8 V/50 W and power supply = 230 V.
17 | Injector Specs
17.1 | Fully automatic microprocessor-controlled contrast medium injector with volume, flow and time control.
17.2 | Wall interface for connection of the contrast medium injector in the examination room, remote from the patient table.
17.3 | Options and accessories
18 | Bracket for Rad.Shield (ER)
18.1 | Accessory bracket for mounting a radiation shield.
18.2 | Radiation Shield
18.3 | StentBoost complete
18.4 | Interventional Tools Hardware
18.5 | Digital subtracted Angio
18.6 | Physio Viewing
18.7 | Rotational Scan
18.8 | Automatic Position Controller
19 | Peripheral and Neuro module
19.1 | Pivot for Table base
19.2 | For angiographic andinterventional procedures of the upper peripherals.
19.3 | Provides improved table access for patient transfer .
19.4 | Allows pivoting of the table base around its vertical axes.
19.5 | Pivot range form -90 degree to + 180 degree ( or-180+90degree) with locked.
19.6 | Positions on o,-13/+13( facilitating arm-angiography) and -90/+90 and 180 degrees.
20 | Clinical Application/Features
20.1 | Real-time frequency subtraction imaging that eliminates the need for mask runs
20.2 | Function of stitching images taken while manually performing table panning in the longitudinal and lateral directions
20.3 | The current position of exposure can be displayed on the stitching image
20.4 | Capable of responding to nonlinear real-time pixel shifts with different movement directions and amounts during DSA acquisition
20.5 | DSA and DA, a harmonization function with dynamic density optimization
20.6 | Transcatheter arterial embolization (TAE) with special FPD application dose reduction
20.7 | DOSE DISPLAY and DOSE RECORD
20.8 | Peripheral/Neuro/Vascular/Stroke/head and neck/Prostate
20.9 | Clinical application of hybrid subtraction in digital fluoroscopy of the vasculatur
20.1 | Embolisation Kit for prostate,Aorto,Endo vascular,AVM,Utrine Etc..
20.11 | Navigation tools, including electromagnetic, optical, laser, and robotic guidance systems, as well as image fusion platforms with machine compatablity.
14 | Haemo-dynamic System with Electrophysiology Monitor
14.1 | Complete Periferal and neuro Lab Physiological Monitoring System offering all of the functionality necessary to monitor patient parameters, automate data collection.
14.2 | should include full cardiac Output measurements with right and left heart values including SVR, PVR.
14.3 | 32 ch Electro-Physiology parameters
14.4 | Comprehensive Hemodynamic analysis
14.5 | Full disclosure
14.6 | Coronary Tree Annotation with Standard Cardiac and Peripheral Templates & Pediatric Congenital Heart Diagrams
14.7 | Based on open architecture like Windows
15 | Additional Workstation & Documentation:
15.1 | Separate workstation with 2 to 4 MP 21” HR Flat panel monitors, DVD writing facility and DICOM software for review of images. The DICOM media [DVD] written using this system should be readable on personal computer with necessary DICOM software coming up automatically.
16 | Accessories and Other Equipment
16.1 | Translucent catheterization arm rest
16.2 | A Movable Single head examination light integrated with the main machine & can be Hand controlled and through foot padel with adjustable foccusing and intensity
16.3 | Two surgical light ceiling mounted with approx.110000 Lux or equivalent with adjustable Magnification and intensity. Able to be fixed and movable for for EP Procedures.
16.4 | Arm rest support extendable Board System with itsTable
16.5 | Intercom facility between control room and exam room
16.6 | Automated Coronary & vascular Power Injector System
16.7 | DICOM connectivity with licenses for DICOM Storage SCU/SCP, Query/Retrieve SCU, Modality work list SCU, Storage Commitment SCU, Basic grayscale Print SCU, & Verification SCU/SCP, MPPS etc.
16.8 | All start-up accessories as per end user to be included
16.9 | Discproducer - DVD/CD printer - color - ink-jetEnabling fast and reliable CD/DVD production with multi tray , features a USB 3.0 SuperSpeed interface, and publishes 15 CDs/30 DVDs per hour (40/12x speed, 0.6/4.7GB data respectively) including printing on the disc surface which provide best-in-class accuracy, and the AcuGripTM robotic arm, which ensures only one disc is grabbed at a time.
17 | Quality and Patient safety
17.1 | Unit must be FDA / CE certified
17.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
17.3 | Unit must comply with Eletrical safety test IEC 60601-1,-2,-18 or equivalent
17.4 | Quality control ,Quality assurance and accaeptance test should be done by the company, QC tools for performance verification done by Enduser, if required.
17.5 | Leaded barriers to fix for the full coverage
17.6 | Safety assessment before pre installation and after installation as per FANR should be done by the company
18 | Accessories
18.A | 1 Three Radiation protection leaded Eye wear 2.Two Panorama Shield 3.Three Leaded Caps 4. Three Leaded LEG wraps 5.Aprons wall mounted 6.Patient positioning set with whole body Gel Pads set and foam positiners set 7.Three Lead Apron with peak west type 8.Four Thyroid coil 9.Three Leaded Cap 10.Lead Gloves 11.Radiation sensors dose dedectors accessories or Third party
Supplied with all standard accessories.
19 | Consumables
Supplied with all standard consumables.
20 | Available Additional Option
| OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
21 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for OT Anesthesia technicians and Anesthesia doctors (End User Training) is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 10 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Continuous Update and Upgrading of the system and all provided software with the system including and limited to DICOM software, during the warranty (any hardware upgrade needed for such software upgrade should be included and provided with no additional cost by the vendor). It is expected from the vendor to always install the latest software versions within 4 months of its official release.
9 | Maximum time needed for parts ordering and replacement (is calculated as a system down time )
10 | Maximum response time for system faults or breakdown calls (max. time for the engineer to be at the site) to be provided.
11 | Warranty must be at laest 3 years comprehensive from date of installation and equipment acceptance.
12 | Vendor Must include Site preparation (pre-Installation) work for proposed model

## MRI Civil & MEP Works
Civil and MEP works for New Interventional radiology system
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Civil and MEP works for Interventional radiology system | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
All Pre-installation requirements should be submitted, including suggested room Design and dimensions .Supplier should offer for comprehensive site preparation/modification & pre-installation work. This include complete pre-installation & site preparation work, such as IT, Electrical, Civil , Mechanical and Medical gasses , required for proper installation and functioning of the offered unit. This include all materials, like IT cables, pulling cables, Data points, connections, switches, interfacing with other modalities or fire alarm, electrical cables, pulling cables, sockets, pipes, drainage, replacing the led shielding for the new/existing room or glass or doors , lighting, HVAC, decoration, wall papers or wall decoration, floor and ceiling plans etc..whenever applicable must be included in the offer as part of the unit. This also includes any site re-designing, drawing, cabinets, doors, A/C installation or re-designing or adding additional A/C, counters and accessories required by the ens-users.
1 | All room preparation and system installation requirements (to be clarified and agreed on after vendor initial site visit)- full room layout required for mini IR angio suite like procedural room,control room,additional storage space,workstation with PACS area,Patient preparatory and recovery area
2 | The location of settings with adequate electrical,oxygen,suction,anesthesia,and emergency services,air exchange and ventilation,lighting, computer terminals,PACSaccess,and monitoring equipment.adequate support area for basic and advanced resuscitation equipment aswell as necessary medications and fluids and additional cooling, to provide lower room temperatures more importantly inside the procedure room
3 | Materials will be used for civil, mep and medical gas should be match the standers of loacal athourtis and UHS (to be clarified and agreed on after vendor initial site visit) |  |  | and
3 | Scope of Proposed Design and As built Drawings ( after completion of Project)Should be Included
4 | Interior Desiging and Decorative works to be Praposed and should Include in scope of Works
5 | Demolition and Mobilization works should be included if reqired dumps should be cleared daily biases or weekly in a way to not effect the health and safety of project also the dumps to municipality is on supplier.
6 | Mobalization(Wall opening and closing ) for cathlab machine delivery should be included
7 | The system offered shall be designed to operate normally under the conditions of UAE, Environmental factors, FNAR and MOH policy the conditions include Power Supply,Climate,Temperature, Humidity,Dust,etc
8 | For any deep drilling or coring, scanning is required.
9 | Cleaning and proper isolation of working area is required, also according to approved execution plan the project maybe divided to phases.
10 | Storage area is not always available at hospital. Inventory storage room should be in close proximity atleast 50square feet ofstorage space to the IR suite The supplier is required to arrange the daily material.
11 | Termination of flooring should be done with proper way to adjacent areas without any level differences.
12 | Efficiency evaluation of the current services system and implement any enhancement or changes if required and according to clear and approved report parallel to UHS Standers.
13 | Detailed Lead test reports for existing lead and compliance certification to be provided at each step to ensure the radiation lekage not only meets but also exceeds the manufacturers requirements.
14 | supply and installation of lead lining Subject to FANR Safety assesment test/Report. The FANR authority to calculate what shielding is needed after safety assesment report by vendor.
15 | To install appropriate internal fit out or lining, lighting and electrical meet international standards.
16 | supply and installation of Warning light and signs
17 | Partition works and Cealing works should be Similar to UHS standards. Proposed Specification need to be provided
18 | Aesthetic Ceiling picture can be done in X-ray room as per client suggestion
19 | Acoustic ceiling tiles.
20 | The Sound insulation can be incorporated into the framework of the walls during assembly and to ensure standard acoustic internal lining.
21 | Wall treatment
22 | To provide clearest observation window with observation glass
23 | The observation Glass shall be anti-reflective coating
24 | Flooring- Medical Grade (similar or better than existing flooring at UHS)
25 | Removing of existing flooring & levelling with self-levelling screed for complete area.
26 | Power Supply Capacity and Specifications
27 | Proposal for Supply and Instation of Electrical works. socket and switches should be via brand and of same type used in UHS premises.(All wires will be concealed and in conduct pipes and the wiring for electrical and network cable will be commensurate with the existing ).
28 | The lighting should be all LED and should support energy conservation
29 | HVAC- Study the current layout and Propose the modifications required
30 | Fire Alarm and Fire Fighting modfication works to be propose according to standerd
31 | Supply and installation Bedhead units for proposed locations including Medical Gas Points and Nurse call system
32 | console area must be equiped with corian tabletop with chairs
33 | Arrangement plans for Intercom system,Ceiling mounted light,Head and arm support,Infusion bottle holder , Arm for radial support and Networking and connectivity for local and remote assess
34 | To have dedicated image-viewing and interpretation area for the interventional radiologist should be in close proximity to IR procedure room , or if not directly inthe IR control room.
35 | To prepare standard space to fix collision avoidance system to check for any conflicts as to their location and placement will arise (with measurement)
36 | The vendor to identifies and seek approval from UHS team regarding construction materials and the intended use of surrounding areas to minimise radiation scatter and shielding required to restrict radiation exposure
37 | To create enough space for staff to safely circulate around equipment,Cabling, wiring, and hoses, all other medical equipment, should be off the floor to eliminate tripping hazards and safety measures.
38 | Special attention to have Doors atlest 4ft minimum width of 1.3 cm, arrangement for Image illuminatorsand suitable space allotments to cover the electrophysiology systems, hemodynamic equipment,anesthesia machines, perfusion systems ,surgical instrument tables and any other guided imaging machine inside the IR procedural area.
39 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
1
Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Equipment. Two Copies each and one Soft Copy for each.
3 | Supplier not allowed to do any hand crafting tasks inside the Cath lab area if the tasks have more pollution rate.
4 | It is mandatory for the Contractor to submit End-of-Sale, End-Of-Support, and End-Of-Life for each individual hardware component - Proof documents from the manufacturer to be attached with the proposal.
5 | Proposal Should be submitted after having site survey and comparison with the current As built
6 | All The works Should Comply with Standards of Civil defence, SEWA and MOH
7 | for any interior design fitting, supplier should maintain 5% of the fixed quantity for maintenance and submitted for client after job compilation
8 | Contractor should discuss the final technical proposal with the technical team before submission. Proposal should include ongoing warranty, support.