## 1
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | ANESTHESIA MACHINE | 3 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
The state -of -the-Art Anesthesia machine ,Advance latest model used for delivering anesthesia agents to the patients during surgery and monitors the vital signs and ventilates the patients not limited to Gas delivery System as Sources: Compatible with German DIN standard hoses with medical-grade oxygen, nitrous oxide Cylinders .
1.1 | Anesthesia machine is used for delivering anesthesia agents to the patients during surgery and monitors the vital signs and ventilates the patients Offered Unit must state of the art system complying with all the local and international patient and environment safety standards and features .
1.2 | Frame/design
1.2.1 | Anesthesia system should be high end three gas system with three gas Oxygen, Nitrous Oxide and Medical Air, double scale flow meter with high and low flow and minimal flow provisions.
1.2.2 | System should be designed such that all components are integrated to minimize dead space
1.2.3 | Should have an independent oxygen flow meter for oxygen delivery and an integrated variable flow suction unit
1.2.4 | Anesthesia machine should have high grade reinforced fiber frame free from oxidation. It should have three drawers. One retractable writing table, and rigid top tray
1.2.5 | System should have at least three drawers and an additional writing surface that can be easily accessed
1.2.6 | Drawers shall all have the ability to lock, and shall be easily removed for the purposes of cleaning and sterilization
1.2.7 | All be displayed on color coded gauges and be visible at all times during operation
1.2.8 | Should have provision to attach two cylinders 1 each for O2 and N2O
1.2.9 | Should have facility of delivering basal flow of oxygen on switching on the machine
1.2.10 | System shall have a second user accessible port for extraction of anesthetic gas when using a non-rebreathing patient circuit. System should also provide the option of returning sample gas to the scavenging system with a dedicated port
1.2.11 | A single pneumatic/electric on/off switch should activate the gas flow and vaporization
1.2.12 | For the ventilator in the event of power loss and should operate for a minimum one hour
1.2.13 | In case power failure, it can shall still be possible to manually ventilate and deliver anesthetic agent
1.2.14 | System should have easily accessible common gas outlet in the event of an emergency and for use of alternate breathing circuits
1.2.15 | Should have unlockable oxygen flush to delivery oxygen flow of approximately 40 l/min
1.2.16 | Should have built-in safety features like O2 failure alarm, N2O cut off, Low O2 pressure, etc.…
1.2.17 | Should have motion sensitive back lighting for vaporizer dial adjustment. Should also have mandatory illumination of the writing table
1.2.18 | The frame should have integrated power outlets to supply a minimum of four external devices
1.2.19 | Should have locking of the front castors by a single central brake mechanism
1.3 | Gas Flow
1.3.1 | The unit shall have a mechanical hypoxic guard system to control the ratio of oxygen and nitrous oxide to ensure a minimum of 25% of oxygen delivery at all times to avoid delivery of hypoxic mixture
1.3.2 | It shall be possible to deliver Air with only basal flow oxygen independent of the above
1.3.3 | Mentioned hypoxic control
1.3.4 | Avoid errors during power failure and electronic malfunction
1.3.5 | Visual display of the gas flow shall be by physical means independent of electrical power
1.3.6 | for all gases to allow suitable solution and accurate control at low total fresh gas flows
1.3.7 | Flow meters should have backlight and antiglare illumination
1.3.8 | Measurement and display of fresh gas flow offering safety for low and minimal flow
1.3.9 | Unit gas hoses must be provided with DINN (GERMAN) fittings /connectors.
1.3.10 | A bag arm with height and positional adjustment shall be available
1.4 | Vaporizers
1.4.1 | Agent Compatibility: Compatible with commonly used inhalation anesthetic agents.
1.4.2 | Agent Delivery: Accurate and Adjustable concentration Delivery of anesthetic agents.
1.4.3 | Temperature Compensation: Ensures consistent vaporization regardless of ambient temperature.
1.4.4 | The unit should accommodate two vaporizers for anesthetic agent delivery to allow easy selection of agent to be used.
1.4.5 | A third vaporizer storage area shall be available as an option
1.5 | Vaporizer should be selectatec type, tool free mounted at will with interlocking facility to allow operation of only one vaporizer a one time
1.5.1 | Vaporizers supplied with the unit shall be routine maintenance free for the life of the product
1.5.2 | Should provide isoflurane and sevoflurane key filled vaporizers
1.5.3 | Breathing system
1.5.4 | All parts of the breathing system that are in contact with patient gas should be latex free and autoclavable. Should not require tools when dismantled for cleaning and sterilization
1.5.5 | Should accept large and small volume absorber canisters
1.4.6 | The ventilator bellows shall be clearly visible and should ascend on expiration to provide a quick visual indicator for system leaks
1.4.7 | Breathing system should have the option of CO2 absorber bypass control that will allow the absorber canisters to be removed without introducing system leaks
1.4.8 | Should have bag/vent selecting valve integrated onto the absorber and should automatically turn on the ventilator when position to vent mode
1.6 | Ventilator Breathing system
Ventilation Modes: Supports both volume-controlled and pressure-controlled Ventilation modes.
Tidal volume Range: Adjustable Tidal volume settings to accommodate various patient sizes.
Respiratory Rate: Adjustable Respiratory Rate settings for controlled breathing.
Inspiratory/Expiratory Ratio: Configurable Ratio to control the duration of inspiration and expiration.
Positive End-Expiratory pressure (PEEP): Adjustable PEEP settings to maintain lung recruitment.
1.6.1 | Ventilator should be pneumatilly driven, electronically controlled and should be ascending bellows type.
1.6.2 | Ventilator should automatically change drive gas should there be a gas depletion
1.6.3 | Ventilator shall have a large color display with touch screen user interface
1.6.4 | Ventilator should have the following ventilation abilities, volume control, decelerating flow, pressure control, SIMV with pressure support and PRVC mode.
1.6.5 | Ventilator should be capable of ventilating diverse range of patient groups from neonates to patients with restrictive airways with tidal volume range between 20ml to 1500ml with single bellows system
1.6.6 | Assisted modes of breathing should be flow triggered
1.6.7 | Ventilator shall have an active proportional exhalation valve to prevent potential of over delivery during pressure modes of ventilation
1.6.8 | Ventilator should have a leak and compliance test that can be done independently of the full system check
1.6.9 | On switching on, the ventilator system should be able to and shall give the user a choice of doing a unit test or bypassing in the case of an emergency
1.6.10 | Ventilator shall compensate for fresh gas flow and compliance of the entire circuit dynamically
1.6.11 | Measurement at the patient end of the circuit(sensor at the patient end) should be provided to compensate for small leakages and compressible volume variability that occur during ventilation
1.6.12 | User should also have the option of setting a pre-set compliance correction where similar circuits are used constantly
1.6.13 | Should provide constant fresh gas flow into the breathing circuit during the inspiratory phase as mandatory
1.6.14 | Ventilator should have the ability to display and store a hospital default as well as individual user preferences for easy selection of ventilation parameters and include screen layout, alarm preferences and ventilation settings
1.6.15 | Users be able to set their own password
1.6.16 | Apnea alarms must be user adjustable to allow for all operating conditions and phases during anesthesia
1.6.17 | Ventilator should have the ability to display and store patient spirometry loops including flow-volume and pressure volume curves
1.6.18 | Ventilator should also display waveforms for flow and airway pressure
1.6.19 | Ventilator shall display measured fresh gas independent of the flow meters
1.6.20 | Ventilator shall display a dynamic compliance measurement and lung mechanics.
1.7 | Integrated Monitoring System
1.7.1 | User Interface:
1.7.2 | Display: Clear and intuitive graphical interface displaying relevant patient and system information.
1.7.3 | Controls: User-friendly controls for adjusting gas flow rates, ventilation parameters, and alarm settings.
1.7.4 | Data Recording: Ability to record and store patient data for future reference or analysis.
1.7.5 | Monitor should be offered with all the necessary modules
1.7.6 | Anesthesia monitoring system should be of modular type and capable of monitoring adult, pediatric and neonatal patients. Should be from same manufacturer as the anesthesia system
1.7.7 | Monitor should have minimum 19" independent flat panel display with multi-color touch screen user interface to ensure all parameters are visible simultaneously
1.7.8 | Modul rack/housing should be independent and shall be able to be placed near to the patient
1.7.9 | Should be capable of 8 traces display. Should have facility to monitor: ECG, NIBP, SpO2, Respiration, Invasive pressure (3) temperatures (2), capnography and bispectral index , NMT. Should have cardiac output port enabled.
1.7.10 | Should have automatic identification and measurement of anesthetic agents, EtCo2, O2 and N2O and MAC value
1.7.11 | Should have depth of anesthesia monitoring using Bispectral index
1.7.12 | Cardiac output monitoring facility using thermo dilution technology with all accessories
1.7.13 | Arrythmia monitoring on all 12 leads
1.7.14 | Inbuilt ST segment analysis and arrythmia detection for all the leads should be available
1.7.15 | Should have hemodynamic, oxygenation and drug dose calculations
1.7.16 | EtCo2 should have both mainstream and side stream in one module
1.7.17 | Respiration should be available with Cardio vascular artifact filter
1.7.18 | OCRG should be available for monitoring neonates
1.7.19 | Alarm parameters should flash red in the presence of high priority alarms(e.g. ventricular fibrillation and systole) and flash yellow in the presence of medium or low priority alarms(e.g. noisy signal, etc.…)
1.7.20 | 24 hours trend data should be displayed
1.7.21 | All monitors including central stations should have similar user interface for easy usage among all clinicians. Modules should be compatible wit transport monitors if required
1.7.22 | Monitors shall provide capability to remote view of real time waveforms via the internet. Should be able to upgrade to software's for electronic flow sheet and full disclosure of all waveforms
1.7.23 | On screen keyboard for entering this data is preferable. Should have USB ports to connect mouse, keyboard, barcode scanner
1.7.24 | Alarm limits status(ON/OFF) must be indicated on-screen for ach parameter and actual parameter settings must be displayed on-screen when alarms are on
1.7.25 | Position of the displayed waveforms and color of the waveforms must be user configurable
1.7.26 | Monitor shall permit the optional ability to receive and display information from other patient devices such as ventilators, infusion pumps and other standalone devices
1.7.27 | All modules should be compatible with all proposed monitors
1.7.28 | Should be supplied with necessary accessories for adult, pediatric and neonatal accessories
1.7.29 | Should be compatible with hospital HIS and should be HL7(Health Level 7) compliant
1.7.30 | Monitor should have capability to accommodate remote viewing of real time waveforms through internet
1.8 | Environmental Factors
1.8.1 | Safe disposal system: AGSS - Anesthetic Gas Scavenging System should be in place
1.8.3 | Shall meet IEC 60601-1-2:2001 (or equivalent)
1 | Trade-in (If relevant)
1.1 | Supplier should offer a Trade-in price for the following existing system:
1.2 | DESCRIPTION MANUFACTURER MODEL YEAR OF PRUCHASE QUANTITIES
1.3 | ANESTHESIA WORK STATION DATEX OHMEDA-GE, USA AISYS 19/07/2011 03
2 | Quality and Patient safety
2.1 | Unit must be FDA / CE certified
2.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
2.3 | Unit must comply with Electrical safety test IEC 60601-1,-2,-18 or equivalent
2.4 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional . Please specify the details clearly and all the requested spec/info must be well supported with reference documents.
2.5 | Unit must be able to integrate with hospital HIS system . For that all the software's and hardware's must be provided by vendor .
3 | Accessories
Supplied with all standard accessories.
4 | Consumables
4.1 | Supplied with all standard consumables.
5 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for OT Anesthesia technicians and Anesthesia doctors (End User Training) is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Warranty must be at least 5 years comprehensive from date of installation and equipment acceptance.