## 1
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| BENCH TOP INCUBATOR (IVF LAB) | 1 | Comply | Supporting Documents Ref. | by Hospital
1 | Specifications |  | Yes/No
1.1 | Capacity: 6 or 10 temperature-controlled chambers. Chamber Volume: 300-350ml
1.2 | Temperature Range: 5 degrees Celsius to 41 degrees Celsius Accuracy Temperature Control: +0.1 degree Celsius
1.3 | CO2 Range: 0% to 20% Accuracy of CO2 range incubator: +0.3%
1.4 | O2 Range: 0.1% to 20% Accuracy of oxygen range incubator: +0.5%
1.5 | Built-in safety features: Analog Output and BMS (Battery management System)
1.6 | Humidification: optional
1.7 | Sensors: Zirconia Dioxide Ceramic O2 Sensor, long life Infra-red sensor and Lid Heaters
1.8 | Optional: A filter that absorbs VOCs (Volatile Organic Compounds) can be connected to the circulation gas line.
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
3.1 | Supplied with all standard consumables.
4 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for OT Anesthesia technicians and Anesthesia doctors (End User Training) is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Warranty must be at least 3 years comprehensive from date of installation and equipment acceptance.