## Ultrasound
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| ULTRASOUND MACHINE FOR RADIOLOGY Department | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Technical offer must be clear in terms of standards , which will be offered with the Ultrasound machine and optional items.Please specify the details clearly and all the requested specs info must be well supported with reference documents.
2 | Premium ultrasound machine capable of top-of-the-line and 4D imaging along with clear imaging for radiology department use and Latest generation electronic phased array system with minimum 64,000 digital processing channels
3 | Higher-end Diagnostic Ultrasound machine whole body application which includes all radiology application not limited to Abdominal, Vascular Peripheral and neurovascular,Obstetrical,Gynecological,Breast,SmallParts,Transcranial,Pediatric and Neonatal, Musculoskeletal,Urological,endocavitary (transvaginal / transrectal) etc. for both adults and paediatric. Latest state of art high end colour doppler system with full digital technology
4 | Broad band convex array transducer with frequency range of I - 6 MHz suitable for general purpose abdominal, obstetrical and gynaecological applications. ➢ Broad band Linear array transducer frequency range of 3 to 10 MHz suitable for Vascular and small parts applications ➢ Broad band Linear array transducer frequency range of 5 to 17 MI-IL suitable for vascular superficial, musculoskeletal, superficial small parts and elastography applications. ➢ 4-D Endocavitary probe - 5-11 MHz or more Endocavitary probe (should have biopsy facility with needle guides. ➢ 4-D volume acquisition transducer of 2 to 6 MHz for 3D and Live 4D imaging for radiology, Obs-gyn and virtual endoscopy applications ➢ Phased array sector / micro-convex transducer I to 5 mHz for cardiac imaging. • Harmonic Imaging should be available in all probes with the following modes and setting for: ➢ Tissue harmonic ➢ Contrast harmonic ➢ Quantification of harmonics imaging • Tissue Harmonic imaging in power doppler imaging mode for improved sensitivity and specificity in differentiating blood/agent from tissue. • Gain control for additional level of flexibility to image quality control. • Real time high frequency 2D for higher resolution and low frequency Doppler for higher sensitivity in all probes. • Frame rate should be 500 FPS or more. • Steerable PW on all phased array probes. • High-definition acoustic zoom for enlarging sections of 2D and color flow images with more acoustic information for greater clarity and detail while maintaining an optimal frame rate. ➢ Modes— 2D, 3D, 4D, B Mode, B/B Mode, M-Mode, steerable PW Doppler, color Doppler, B/M Mode, B/PW Doppler, B/Power Angio, B and Power Angio should be available. System shall have curved and endovaginal 4D capabilities. • The system shall also support the following: ➢ tissue aberration correction capabilities that enhance clinical performance for technically challenging and obese patients and in breast imaging applications. ➢ Contrast Specific imaging with 2DHarmonics using Pulse Inversion Technology ➢ Panoramic real time compounding imaging 3D RTC imaging with speckle reduction ➢ 4D RTC imaging with speckle reduction ➢ Micro Fine 2D grayscale imaging ➢ Chroma imaging ➢ High PRF Doppler ➢ Simultaneous 2D, color Doppler and PW Doppler (Triple Mode) ➢ extended field-of-view capabilities of continuous-field imaging of at least three times the transducer's active aperture. ➢ Real-time micro-calcification detection technology for Breast application • The system shall support full screen display of all 3D views including individual X, Y and Z MPR views and simultaneous display of thumbnail views on the same system display monitor. • The system shall support display of all multiplanar views and the rendered image during 4D acquisition. The system shall support volume measurements and analysis on quantitative 3D and 4D data with live 4D editing. • The system shall support simultaneous display of volume and multiplanar (MPR) views. • All advanced 3D /4D software applications for radiology, ob-gyn and breast to be quoted. • Monitor should be High resolution medical grade, non-interlaced LCD Color monitor of 19 inches or more with tilt and swivel facility to view in all angles and all light conditions • Color flow imaging for ➢ Increased lateral & spatial resolution. Color flow with capability of automatically picking up color flow as a function of focal depth. ➢ Detection of even subtle areas of turbulence, displaying a more physiological blood flow appearance without loss of frame rate. ➢ Tissue colorization (B-Color) for improved contrast resolution. • Advanced Contrast Imaging should be standard on the system. Convex, linear and endocavitary probes should be capable of contrast imaging, please confirm this in the technical bid. Should have Low Mechanical Index (MI) and Flash Modes and start/stop timer. Real-time display of predicted Thermal Index (TI) and predicted Mechanical Index (MI) values should be provided in Contrast Imaging All advanced software package with wash in, stay and wash out with contrast quantification package should be available. • Should have facilities and extensive application software for adult abdominal, Obs/gyn, peripheral vascular, musculoskeletal, small parts imaging and endocavitary applications. (All application package should be built into the system). • Cine loop facility, both frame by frame and in cine mode, with a memory for atleast 2000 2D color images' review and atleast 30 seconds of doppler and M mode data. • High frame rate review for better clarity of play back images study in slow motion. • Quad loop with memory for pre and post image comparison of any procedure. • Memory —2000 frames or more in quad loop. M Mode & Doppler scroll memory -40 seconds or more. • Frame grabber facility for post analysis. Various maps for pre and post processing. ELASTOGRAPHY • System shall be equipped to perform elasticity imaging using latest available technology in a variety of application (liver, breast, prostate) and on a variety of transducers (convex, linear and endo-cavitary) accompanied by necessary quantification package software. • One touch entry into elastography mode. Elastogram applied as a region of interest box with user control of size and location through entire field of view Realtime indicator for elastogram quality. Single screen 2D with elastogram and side-by-side display of 2D image and 2D with elastogram. • Shadow duplication and measurement capability in side-by-side display. • Distance and area tools. Acquisition rate control for elastogram optimization. Elastogram optimization settings to optimize to different tissue compositions. •
5 | System Dynamic Range should more than 160db. • User defined system and application pre-sets for multi-user department. The number of application pre-sets is to be mentioned. • In- built hard disk storage capacity of atleast 160 GB with facility of direct storage and retrieval of B/W and color images (both frozen and cine loops). CD, DVD drive for read and write of stored images. • Depth of Field of 30 cm or above should be available • PRF Range should be 500 Hz to 50,000 Hz or higher. • Alpha numeric key board with illuminated keys and status display. Key panel Height Adjustment Should be Possible. All panel key should be customized, including Freeze Key. • Color Map resolution up to 256 levels. • Facility for high definition digital acquisition, review and editing for complete patient studies. • Unit should have 4 transducers holders and one gel bottle holder. • 3 Active Ports should be available. 2 parking ports. Any Probes any Port interchangeable connectivity should be possible with simple electronic selection method for interchanging transducers. ➢ Detailed Radiology, Obs &gyn and vascular measurement packages should be available. ➢ System should have extensive calculation packages. ➢ Distance, volume, Area, % stenosis on B mode. ➢ Distance, Time, Heart Rate, Slope on M mode. ➢ Velocity, Acceleration time, Slope, P1, RI, S/D Ratio with Auto Doppler calculation on Doppler mode • The system should have Up I Down & Right I Left Image rotation, One touch Image optimization and Edge Enhancement settings. ➢ The system shall be capable of supporting color Doppler imaging on all phased, linear, motorized 3D, and curved array transducers. ➢ Color Power Angio imaging which enhances visualization of blood flow in very small vessels and tissue vascular beds shall be supported. ➢ Spectrum Imaging: PW Doppler Mode should be available. ➢ Doppler sample volume size shall be adjustable from 0.5-20 mm. ➢ The system should have gate adjustments on spectral modes, auto angle correction, filter adjustment, base line and sweep speed adjustments. The system shall have adjustable wall filter settings from 40-300 Hz. ➢ Cineloop Review and Doppler Review must be independently selectable. Automated, real-time analysis of Doppler spectral waveform which is available in live or frozen modes. ➢ Real time panoramic imaging to have an extended field of view of structures. ➢ System should have facility for separate 2D quick scan (auto 2D optimization)I Doppler Quick Scan (auto baseline and PRF adjustment). ➢ System should have automatic real time quantification of doppler parameters. • Virtual Convex (Trapezoid) format with both Linear as well as convex Probes should be available
6 | Transducer Types: Should have four (4) Active Probe slots. • Convex Array Min 1-6 Mhz • Linear Array Min 3-10 Mhz • Matrix Array Min 6-15 Mhz • Small Linear probe Min L8-18 Mhz • Trans Vaginal/Rectal Probe
7 | Contrast-Enhanced Ultrasound with TIC quantification
8 | Volume Navigation
9 | Operating Modes • B-Mode • M-Mode • Color Flow Mode (CFM) • Extended Field of View • Power Doppler Imaging (PDI) • PW Doppler • CW Doppler • Volume Modes (3D/4D): • 3D Static • 4D Real Time • Anatomical M-Mode • Coded Contrast Imaging (option) • Elastography •B-flow
10 | Unit should be fully supported to System should be DICOM 3 or higher version compatible and capable of being interfaced with HIS/RIS/PACS and connectivity to any PC/computer etc in DICOM format.
11 | •Monitor Screen minimum 19′′ High-Resolution LCD •Min 10′′High Resolution Touch LCD screen for software menu
12 | Specity Min/Max hight adjustment ,Speckle reduction filter, real time spatial compounding, frequency compounding or latest technology should be available in all probes for better resolution and penetration. The system shall have automatic system optimization (One Button) for Both B Mode and Doppler
13 | Horizontal/Vertical viewing angle of monitor
14 | Image Storage: •Storage Formats:DICOM – compressed/uncompressed, single/multiframe with/ without RawData •Export JPEG, WMV (MPEG 4), and AVI formats • support Storage Devices:USB Memory Stick,CD,DVD, Hard Drive stoarge
15 | Peripheral Options • User Interface Operator Keyboard • Digital B&W thermal printer • DVD video recorder • Digital color thermal printer • Foot Switch with programmable functionality
16 | Connectivity & DICOM • Ethernet network connection• Verify• Print• Store• Modality Worklist • Storage Commitment• Query / Retrieve
17 | Having B/M color and PW/CW and 512 grey shades or more for sharp contrast resolution
18 | Unit Must haveHDMI /DVI/VGA ports
19 | Frame rate/Sampling rate must not be less then 150 frames/sec
20 | Must be supplied with 500GB HDD or more
21 | Full DICOM and HL7 compatability and All new software should be updated free of cost for at least 5 years.
22 | Specify Dimensions andweigjt of the uint
23 | Unit should be provided with all necessary accessories and consumables.
24 | Must have universal power adapter 220-240v ,50/60HZ
Quality and Patient safety
25 | Unit must be FDA / CE marked .
26 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
27 | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1..
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and for Nurses and Physicians (End User/Application Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .