## Sheet2
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | ALLERGY TESTING ANALYZER | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
2 | Specify Make/Model/Origin
3 | Size: medium
4 | Easy to use
5 | Connection to track care
6 | Software: Ability to view the device from a software and review of QC & L-J charts
7 | Maintenance: Auto-maintenance.
8 | Quality control: Manufacturer Quality control availability
9 | System Capability: Quantitative measurement of specific IgE, total IgE, allergen panels.
10 | Sample Handling: Fully automated loading, dilution, incubation, washing, reading.
11 | Allergen Coverage: Food, inhalant, environmental, dust mites, molds, pet dander, region-specific allergens.
12 | Calibration: Multi‑point calibration with traceability.
13 | Certification: FDA or CE certified
14 | External Quality control: Listed in CAP PT
15 | Automatic barcode identification of samples
16 | User friendly, Easy loading and unloading of samples
17 | Unit must be antimicrobial Protection-Easy to clean surface
18 | Specify Weight and Dimensions of the unit
19 | Power consumption & Heat dissipation in BTU/h
20 | Unit must have LIS integration
21 | Interface Requirements for HIS Integration (Trackcare-TCLE)
22 | Protocol Support The machine must support ASTM or HL7 protocols to enable seamless integration with the Laboratory Information System (LIS)/TrakCare Lab Enterprise.
23 | Software Components Any required software components, middleware, or workstation setups necessary for integration must be provided by the supplier.
24 | Skilled Resources The supplier should have trained and skilled personnel with expertise in LIS/HIS integration to support implementation and troubleshooting.
25 | Documentation
26 | The supplier is required to provide detailed LIS integration specification documents covering all technical aspects of the interface.
27 | Unit must be FDA / CE certified
28 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
29 | Unit must comply with Eletrical safety test IEC 60601-1,-2,-18 or equivalent
30 | User-friendly operation and smooth movement - One-unit package with in-built battery/generator system - Durable with operational implication for different position and parts of the body.
31 | supplier must provide the information on energy efficiency options in the equipment, sustainability methods adopted, any carbon footprint technologies utilized, and certifications
32 | Accessories
Supplied with all standard accessories.
33 | Consumables
Supplied with all standard consumables.
34 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
24 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for OT Anesthesia technicians and Anesthesia doctors (End User Training) is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Continuous Update and Upgrading of the system and all provided software with the system including and limited to DICOM software, during the warranty (any hardware upgrade needed for such software upgrade should be included and provided with no additional cost by the vendor). It is expected from the vendor to always install the latest software versions within 4 months of its official release.
9 | Maximum time needed for parts ordering and replacement (is calculated as a system down time )
10 | Maximum response time for system faults or breakdown calls (max. time for the engineer to be at the site) to be provided.
11 | Warranty must be at least 5 years comprehensive from date of installation and equipment acceptance.