## IOL Machine
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | IOL (Intraocular Lens) swept-source optical biometry | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
2 | Please specify make, model, and origin of the machine.
3 | The device should be capable of accurately measuring the following parameters:
4 | Axial Length (AL): Accuracy of ±0.01 mm.
5 | Corneal Curvature: Measurement range of 35–55 D.
6 | Anterior Chamber Depth (ACD): Accuracy of ±0.1 mm.
7 | Lens Thickness: Measurement accuracy of ±0.1 mm.
8 | Pupil Diameter: Measurement range of 2.0–9.0 mm.
9 | Keratometry (K): Accurate to ±0.05 D.
10 | Swept-Source Technology: The system must use swept-source OCT (optical coherence tomography) with a wavelength of at least 1050 nm for deeper penetration into ocular structures.
11 | Measurement Speed: Fast measurement time, with an ideal value of less than 1 second per eye.
12 | Refractive Index Compensation: Should provide real-time adjustments for different refractive indices of the cornea and lens.
13 | IOL Power Calculation: The system must be capable of calculating IOL power using different IOL formulas (e.g., SRK-T, Holladay, Hoffer Q, Barrett).
14 | IOL Compatibility: The system must be compatible with a wide range of IOL models and manufacturers.
15 | IOL Design Options: Ability to support standard monofocal, toric, multifocal, and extended depth-of-focus (EDOF) IOLs.
16 | Automatic IOL Power Selection: The system should provide automated IOL power suggestions based on biometric data.
17 | Measurement Error: The device should demonstrate a measurement error of less than ±0.02 mm for axial length and less than ±0.05 mm for keratometry.
18 | The system should have high reproducibility for multiple measurements taken under similar conditions, with minimal variation.
19 | Screen Display: The system must have a user-friendly touch-screen interface with clear display of measurement results.
20 | Software: The system should come with advanced software for data analysis, storage, and printing of reports. The software should also allow for exporting data to other HIS systems
21 | The system should feature automatic calibration or be easily calibrated using external devices.
22 | Must provide biometric measurements (axial length, corneal curvature, etc.) for both eyes in less than 30 seconds per eye.
23 | The device should feature non-contact measurements with no risk of eye injury or discomfort during the process.
24 | The device must provide clinical-grade accuracy validated by peer-reviewed studies and real-world performance data
25 | Unit must be antomicrobial protection
26 | Power supply: 220–240V, 50/60 Hz
27 | Quality and Patient safety
27.1 | Device should be CE and FDA approved
27.2 | Please specify the weight and unit must be provided with good quality castors .
27.3 | Offered System must comply with Electrical safety IEC 6060-1
27.4 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
28 | Accessories
Supplied with all standard accessories -Handpiece with tips,cutters.
29 | Consumables
Supplied with all standard consumables-Starter kit of disposables,Light probes infusion cannuls
30 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
31 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .

## Opthalmic Units and chair
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Opthalmic Unit with Patient Chair | 2 | Comply | Supporting Documents Ref. | by Hospital
2 | opthalmic surgery Chair | 1
| Specifications |  | Yes/No
1 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
2 | Ophthalmic Examination Unit with integrated patient chair.
3 | Please specify make, model, and origin of the machine.
4 | The unit must be modular,Adjustable (Up/Down) ,sliding tabletop designed for use with various diagnostic and therapeutic devices (e.g. slit lamp, autorefractor, tonometer, fundus camera).
5 | The unit must include adjustable mounts for key ophthalmic instruments.
6 | Must support smooth height adjustment for different operators, with at least 180° of rotation to accommodate both left- and right-handed users.
7 | A drawers or platform for placing instruments, with sufficient weight capacity and smooth sliding options.
8 | Integrated lighting system with adjustable brightness and color temperature for optimal illumination during eye examinations.
9 | LED or halogen lighting with no flicker.
10 | Variable intensity adjustment to meet various diagnostic needs.
11 | The unit will be installed in an examination room with standard clinic dimensions(To be quoted once after site visit )
12 | Patient Chair
12.1 | The patient chair should have a height adjustment range of at least 45-85 cm to allow easy access for different height patients and clinicians.
12.2 | The backrest should recline smoothly with an adjustable range of at least 0° to 90° for optimal patient positioning.
12.3 | The footrest should be adjustable and should support different leg lengths and patient positions.
12.4 | Adjustable armrests for patient comfort and stability during the exam.
12.5 | Adjustable and padded headrest to ensure patient comfort and support during examinations.
12.6 | High-quality, soft, non-toxic padding for patient comfort. The materials should be easy to clean and maintain.
12.7 | The chair should support a minimum weight of 200 kg
13 | Please specify make, model, and origin of the machine.
14 | Height of the chair should be easily adjustable
15 | Chair should be rotatable 360 degrees smoothly
16 | opthalmic surgery Chair | 1
16.1 | The chair should support a minimum weight of 150 kg
16.2 | Chair should be rotatable 360 degrees smoothly
17 | Quality and Patient safety
17.1 | Device should be CE and FDA approved
17.2 | System should not be bulky , Please specify the weight and unit must be provided with good quality trolly .
17.3 | Offered System must comply with Electrical safety IEC 6060-1
17.4 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
18 | Accessories
Supplied with all standard accessories.
19 | Consumables
Supplied with all standard consumables.
20 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 10 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .

## Phacoemulsification -Vitrectomy
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | Phacoemulsification -Vitrectomy machine | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
1 | Please specify make, model, and origin of the machine.
2 | phacoemulsification and vitrectomy system-fully able to support both anterior and posterior surgeries in one platform.
3 | Phacoemulsification
3.1 | Ultrasound Technology:
a | Torsional and/or longitudinal ultrasound
b | Adjustable frequency and power
c | Pulse, burst, and continuous modes
d | Real-time energy modulation
3.2 | Fluidics System-Adaptive fluidics or gravity-based fluidics with electronic control
a | Vacuum range: minimum 0–600 mmHg
b | Precise intraocular pressure (IOP) control
c | Flow rate: minimum up to 50 to 60 cc/min
d | Anti-surge protection system
3.3 | Handpiece-Lightweight, autoclavable,Compatible with micro-incision cataract surgery
a | Cassette / Tubing-Single-use disposable cassette
4 | Vitrectomy Unit
4.1 | Cutting System
a | Electric or pneumatic cutter
b | Cutting speed: minimum 5,000–10,000 cuts per minute (or higher)
c | Adjustable duty cycle
d | Modes:Core, shave, and proportional modes
e | Anterior and posterior vitrectomy capability
4.2 | Aspiration & Vacuum
a | Vacuum range: up to at least 650 mmHg
b | High-flow aspiration with surge protection
4.3 | Gauge Compatibility-Compatible with 20G, 23G, 25G, and 27G systems
4.4 | Cassette / Tubing-Single-use disposable cassette,Easy installation with automatic recognition
5 | Illumination System
a | High-intensity LED or Xenon light source with filters
b | Adjustable brightness
c | Compatible with standard fiber-optic probes
d | Built-in safety filter to reduce retinal phototoxicity
6 | User Interface & Control
a | Large color touchscreen display,Graphical display of real-time parameters
b | Customizable surgeon profiles
c | Foot pedal with:Dual Linear control,Multifunction capability,Wireless option
7 | Unit must be antomicrobial protection
8 | Power supply: 220–240V, 50/60 Hz
9 | Quality and Patient safety
9.1 | Device should be CE and FDA approved
9.2 | Please specify the weight and unit must be provided with good quality castors .
9.3 | Offered System must comply with Electrical safety IEC 6060-1
9.4 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
10 | Accessories
10.1 | Supplied with all standard accessories -Handpiece with tips,cutters.
11 | Consumables
11.1 | Supplied with all standard consumables-Starter kit of disposables,Light probes infusion cannuls
13 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
14 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .

## Retinal cryotherapy machine
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Retinal cryotherapy machine | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
2 | Please specify make, model, and origin of the machine.
3 | Retinal cryotherapy machines and probes are designed for use in ophthalmic surgeries, including cryopexy for retinal detachment, cyclodestructive procedures for refractory glaucoma, vitreous fragment extraction, cataract surgery, cryodestruction of lash follicles for trichiasis, and treatment of retinopathy of prematurity (ROP)
4 | The machine should use either a nitrous oxide (N2O) or compressed gas system for cryogenic freezing.
5 | The machine should be able to control the temperature range of the cryoprobe with precision (typically between -50°C and -90°C).
6 | The cryoprobe must be ergonomically designed to allow for fine control and access to difficult-to-reach retinal areas.
7 | The machine should offer different cryo tip sizes for varied clinical situations.
8 | The machine should have an automatic freeze-thaw cycle with adjustable freeze time to avoid excessive tissue damage.
9 | It must include mechanisms to avoid over-freezing, ensuring patient safety.
10 | The system should maintain temperature with a tolerance of ± 1°C.
11 | The machine should have adjustable freeze and thaw cycle times ranging from 1–5 minutes.
12 | The machine should have adjustable pressure settings with real-time monitoring to ensure consistency.
13 | The machine should have a clear, easy-to-read digital interface with real-time feedback on the treatment process.
14 | Should allow the operator to adjust cryotherapy parameters such as temperature, time, and pressure.
15 | The device should be able to display information on an external monitor for enhanced visualization .
16 | The system should provide audio cues or visual indicators for critical stages of the treatment .
17 | Unit must be antomicrobial protection
18 | Power supply: 220–240V, 50/60 Hz
19 | Quality and Patient safety
19.1 | Device should be CE and FDA approved
19.2 | Please specify the weight and unit must be provided with good quality castors .
19.3 | Offered System must comply with Electrical safety IEC 6060-1
19.4 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
20 | Accessories
Supplied with all standard accessories -Standerd set of retinal probes (reusable and disposable)
21 | Consumables
Supplied with all standard consumables
22 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
23 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Opthalmology doctors and other associated staff and Service training for Biomedical Engineeing Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .