## ENT Workstation
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | High End ENT unit without Microscope | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
Technical offer must be clear in terms of standards , which will be offered with the ENT machine and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
1 | Medical Grade Steel housing:550 mm with PUR shaped top parts, Easy access to air filter, condensation discharge and liquid bottle.
2 | Large instrument surface: Two storied, with aluminium trays and set of stainless steel dividers and plexiglass cover, to protect instruments.
Aluminium trays and set of stainless steel dividers
3 | Compressed air unit: continuously adjustable from 0,1-4 bar,(spray/politzer), fine regulation handpiece pressure limiter.
4 | Preferably 3 Spray bottles in different colours: front cannulas in two different diameters straight.Preferably System must contain minimum of 1,5 Ltr. tank for compressed air, so that the compressor is required to work only 4-5 times or less per day.The compressor must be completly separated from the suction unit, silently motorized. The chrome plated steel handle is adjustable to politzer olives.A set of politzer olives must be included.
5 | Motor suction unit: must be silently motorized, 0,92 bar max.vacuum, 40-55 Ltr./min., bacterial filter, double overflow control fuse, 1,5 Ltr. liquid container.The suction tube must be equipped with an automatic on/off switch for the vacuum function.An ear rinsing funnel must be included in the System.
6 | Drawer as a writing surface:with removable plexiglass shelf insert / or paper dispenser drawer.
7 | Instrument drawer: with aluminium trays and set stainless steel dividers.
8 | It must contain Empty compartment with plug for additional biomedical equipment.
9 | Kidney shaped ear rinsing funnel:on a swivel support including ear wax separator with additional suction tube connected to the liquid Jar and it must be operated through a switch in the front panel.
support including ear wax separator .With
additional suction tube connected to the liquid jar.
must be operated through a switch in the front panel.
10 | Must equiipped with Warm water rinsing device: complete with chrome plated handle, handle must be autoclavable, snap closure system, to connect the tube, fine spray regulationvalve, 38° water temperature, water is heated in a separate tank system, to prevent mineral built-up,as several front cannulas (Luer) and a water supply hose. Including autoclavable water filter with cartridge.
11 | Suction & Irrigation device: It should be offered with autoclavable and steel handle which incorporates two tubes that are connected to suction and irrigation.The distal part of the steel handle adapts to a tonsil shaped glass tub with glass pipe that can easily be exchanged after the treatment. The basic set must be offered and equipped with a set of autoclavable glass pipes.
12 | Sinus Rinsing handle:
Suction / Irrigation handle must be offered with two nasal olives and in three different sizes S-M-L.
Intermediate nostril distance adjustable to various nasal shapes via gooseneck
One side irrigates while the other side simultaneously evacuates with connection to the hose water suply.
13 | Suction pipe cleaner: Preferably availble with exchangeable, reusable, chrome plated adapter shaped like a cone, to make it suitable for tubes and suction pipes of different diameters.The reservoir for the disinfection liquid (1 Ltr.) is located in the service box. Adapter is autoclavable.
14 | Wall or floor connection box: with control valve, to connect the water supply and water drain hose
15 | Endoscopy center consists of:
a | Cold light source minumim 4 outlets: Preferably LED outles
b | 3 warming quivers for rigid endoscopes, 38° C,thermostat control with removable disinfectable insert and botton screw for easy cleaning .
c | 2 quivers for flexible endoscopes with removable insert.
d | 3 disinfection quivers for rigid endoscopes with removable insert .
e | Automatic on/off switch: of the single light outlets with light barrier and also with disinfection time control .
16 | Additional shelf is required to place additional medical equipments .
17 | Instrument cabinet must be Stainless stell and having large surface ;
a | Steel housing 550 mm with PUR shaped top parts,Large instrument surface with aluminium trays,set of stainless steel dividers and plexiglass dust cover
b | Large instrument surface with aluminium trays and it must contains minimum as;
c | Compartment for bottles behind cover.
d | 1 drawer for storage.
e | 1 instrument drawer with aluminium trays and set of stainless steel dividers.
f | 1 empty compartment with plug for additional equipment.
g | 1 writing surface
18 | Examination Chair
a | Motor support, integrated foot switch ; Stroke 20 cm; adjustable from 51 - 71 cm ;motor support for simultaneous movement of backrest and footrest ; backrest can be inclined to horizontal position ; Rotation 180° left and 180° ; fully synthetic leather cover black ; Mobile on castors ;Adjustable Headrest and Remotely controlled ( optional)
19 | Doctor’s Chair
a | Available with adjustment of . backrest & adjust. seat height minimum of 50 to 80 cm
b | Must be Having Seat adjustment of min 40 cm diameter, mobile on 5 safety castors for good support .
c | Steel parts chrome platedsynthetic leather cover
20 | Endoscopy Camera HD ( Please mention the techical details in your technical offer in detail and it must be supported by catalogues and brochures)
a | It should be prvided with High definition Endoscopy and microscopy camera ; Preferably having Resolution 1920x1200 pixels Preferably equipped with Globe shutter technology ; With minimum 15.6” Touchscreen 45 ° and with Swivel mount ; Must equipped with C-Mos chip ( allows the maximum amplification of the remaining light) and must be having option to connect second camera head.Preferably having SSD Data storage 128GB for up to 15 hours of Video recording Option must be available for end user to choose between USB or network transfer . System must be HL7 and DICOM interface for integration to our hopsital HIS. operation via screen keys or external keyboard ; outlets for further screens, footswitch .
21 | Additional monitor 32" is required ( optional)
22 | Diagnostic Package includes ;
a | Rigid Endoscope ~ Ø 4 mm, 175 mm length, 0°
b | Rigid Endoscope ~ Ø 4 mm, 175 mm length, 30°
c | Head Light ( Please specify clearly the brand / manufacturer and specs)
d | with light transmission cord
e | Light Cord 180 cm length, Storz fitting
f | Tele-Laryngo-Pharyngoscope HD L90° F adjustable focus
24 | LED Headlight
a | 1 W super-high output LED,
b | rated at minumim of 50,000 hours operating life,
c | mPak and preferably Li-ion battery and with plug in adapter .
25 | System must be having Electrical safety test IEC 6060-1
26 | Device must be manufactured on good standards like Qulaity management system , ISO-13485
27 | Offered machine must be having good safety features , equipped with safety alrams , technical alrams and system interlocks.
28 | Offered unit -History of incident,recalls and safety alerts .
29 | ENT INSTRUMENT SET (OPTIONAL)
Ear funnel size 3
Ear funnel size 4
Ear funnel size 5
Ear funnel size 6
Ear funnel size 7.5
Jansen forceps 16.5 mm
Nasal speculum 75 mm
Nasal speculum for children
Nasal speculum for adults
Ear forceps 8 cm
Ear forceps 6 cm
Cotton applicator 15 cm
Cotton applicator 13 cm
T.C. needle holder 15 cm
Suction pipes 3 mm
Suction pipes 3.5 mm
Suction pipes 4.5 mm
Ear curettes
Ear hooks, 14.5 cm
Drainage pipes 2.5 mm
Nasal dressing forceps 6 cm
Suction pipes 10 cm, dia 1.5 mm
Suction pipes 10 cm, dia 2 mm
Suction pipes 10 cm, dia 2.5 mm
Adapter for suction pipes
Bowman probe 00-0
Bowman probe 1-2
Joseph scissors 14 cm
Surgical handle 12.5 cm
Blades No.11
Blades No.15
Lichtwitz trocar 10 cm
Kidney tray 25 cm
Scissors straight 14.5 cm
Dressing forceps 14.5 cm
Forceps serrated 20 cm
Tongue depressor Brünings 19 cm
Parazentese needle bajonett
Lucae forceps 14.5 cm
Tissue forceps 1 x 2 teeth 14.5 cm
30 | Accessories
Supplied with all standard accessories.
31 | Consumables
Supplied with all standard consumables.
32 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
33 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and for Nurses and ENT Physicians (End User Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .