## Air pendrive drill system
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To Be Filled
| Electric Pendrive drill system | 1 | Comply | Supporting Documents Reference | By Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
2 | Spacify Origin,Make and Model of the unit
3 | Offered model should not have a recall, and in case the offered model is purchased and there is a recall during warranty then the supplier has to provide a replcement within 5 days with the same features of the recalled model.
4 | The offered models of each part of the complete system should be from same manufacturer brand and Should have an autorized distributor in UAE.
5 | Able to integrated in future with the Specific same brand Navigation System.
7 | The Required system should have the below specifications.
8 | Console unit
9 | Electric pen drive Handpiece
10 | Variable speed settings from low RPMs for delicate bone work to over 35,000 RPM for high-speed cutting
11 | Micro motor complete with motor cord not less than 3 m in length
12 | Console shall be Touch Screen & Compact
13 | Low noise level- quiet operation, the system minimizes acoustic disruption in the operating room
14 | Supports both forward and reverse rotation, selectable via the console or footswitch, facilitating safer burr removal and enhancing procedural versatility
15 | Variable and fixed speed setting that can be switched in the console and the foot switch.
16 | Multifunction foot pedal
17 | Controls to switch between variable and fixed speed, reverse and forward rotation.
18 | Hoop locks firmly in place, then collapses easily for storage.
19 | Non-slip design and compact size.
20 | Foot control for motor operation.
21 | Motor Type
22 | High-speed electric motor delivers consistent torque and maintains speed under load, supporting high-performance outcomes.
23 | Able to be integrated with navigation system enables the surgeon to align burr trajectory with preoperative planning, improving surgical confidence and reducing risks.
24 | Available with an optional finger-controlled handpiece, allowing surgeons direct control of drill activation and speed adjustment without relying on a footswitch
25 | Consumables
26 | Specific consumables for different length attachment, offering tailored performance and optimized efficiency.
27 | Attachment & Functional Accessories:
27.1 | Electric pen drive Handpiece with cable
27.2 | Drill attachment with mini coupling
27.3 | AO/Asif Drill Attachment
27.4 | Oscillating and Sagittal Saw Attachments
27.5 | Reciprocating Saw Attachment
27.6 | Burr Attachment
27.7 | Irrigation Nozzles
27.8 | Instruments tray with lid
28 | Unit must be antimicrobial Protection and water resistance
29 | Specify Dimensions and weight of the units
30 | Unit should be provided with all necessary accessories and consumables.
31 | Must have universal power adapter 220-240v ,50/60HZ
32 | Quality and Patient safety
32.1 | Unit must be FDA / CE certified
32.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
32.3 | Unit must comply with Electrical safety test IEC 60601-1,-2,-18 or equivalent
32+A41:A47.4 | Quality control ,Quality assurance and acceptance test should be done by the company, QC tools for performance verification done by End-user, if required.
32.5 | Safety assessment should be done by the company
33 | Consumables
34 | Vendor must include prices of all standard and optional consumables.
34 | Available Additional Options
No. | Options Description |  |  |  | REQUIRED OR NOT REQUIRED
36 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for Nurses and Physicians (End User Training) is mandatory.
3 | IEC 60601-2-38 Medical Electrical Equipment's part 2-38 Particular requirements for safety of Electrically Operated Hospital Beds
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Should have local service facility .The service provider should have the necessary equipment's recommended by the manufacturer to carry out preventive maintenance test as per guidelines provided in the service/maintenance manual
7 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
8 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
9 | Warranty : Comprehensive at least 3 Years
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Should have local service facility .The service provider should have the necessary equipment's recommended by the manufacturer to carry out preventive maintenance test as per guidelines provided in the service/maintenance manual
7 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
8 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
9 | Warranty : Comprehensive at least 3 Years