## Neuro-Spine drill system
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To Be Filled
| Neuro-spine surgery drill system | 1 | Comply | Supporting Documents Reference | By Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
2 | Spacify Origin,Make and Model of the unit
3 | Offered model should not have a recall, and in case the offered model is purchased and there is a recall during warranty then the supplier has to provide a replcement within 5 days with the same features of the recalled model.
4 | The offered models of each part of the complete system should be from same manufacturer brand and Should have an autorized distributor in UAE.
5 | Able to integrated in future with the Specific same brand Navigation System.
6 | System to be complete with in pump for irrigation.
7 | The Required system should have the below specifications.
Console
8 | Remote control for irrigation ensuring optimal cooling and debris clearance during high-speed drilling procedures
9 | Variable speed settings from low RPMs for delicate bone work to over 75,000 RPM for high-speed cutting
10 | Micro motor complete with motor cord not less than 3 m in length | 2
11 | Console shall be Touch Screen & Compact
12 | Low noise level- quiet operation, the system minimizes acoustic disruption in the operating room
13 | Forward and reverse rotation option. Supports both forward and reverse rotation, selectable via the console or footswitch, facilitating safer burr removal and enhancing procedural versatility
14 | Variable and fixed speed setting that can be switched in the console and the foot switch.
15 | Multifunction foot pedal
16 | Controls to switch between variable and fixed speed, reverse and forward rotation.
17 | Hoop locks firmly in place, then collapses easily for storage.
18 | Non-slip design and compact size.
19 | Foot control for motor operation.
20 | Motor Type
20.1 | High-speed electric motor delivers consistent torque and maintains speed under load, supporting high-performance outcomes.
20.2 | Able to be integrated with navigation system enables the surgeon to align burr trajectory with preoperative planning, improving surgical confidence and reducing risks.
20.4 | Available with an optional finger-controlled handpiece, allowing surgeons direct control of drill activation and speed adjustment without relying on a footswitch
20.5 | Consumables
22 | Specific consumables for different length attachment, offering tailored performance and optimized efficiency.
22.1 | A comprehensive set of specialized drill bits and burrs is available, supporting diverse clinical procedures including: • Spine Surgeries • Lateral Skull Base Access (e.g., translabyrinthine, retrosigmoid) • Transnasal and Endoscopic Approaches to the skull base • Posterior Cervical Decompression and other spinal access techniques • These burrs are designed for optimal reach, cutting precision, and compatibility with navigation-guided procedures.
22.2 | The system offers a complete anterior skull base surgical solution, integrating low-profile handpieces, curved burrs, and high-speed operation to enhance visualization in narrow operative corridors
22.3 | Curved burs allow better visibility and access. Integrated irrigation option for better visibility and performance.
22.4 | Craniotome consumable with dura guard for adult and pediatric.
23 | Attachment Types(Neuro-spine):
23.1 | Straight and angled attachments: 1)Foot rotating F1 B-1 2.4 MM 2)Foot rotating F2 B-5 2.4 MM 3)7 CM Stright 2.4 MM Spine surgery 1)10 CM angled 2.4 MM 2)15 CM Angled 2.4 MM 3)Stright & angled attchment 7.5 cm
24 | Variable exposure attachments allow you to vary dissecting tool exposure. Small bore variable attachments can increase exposure up to 10mm in 10 increments and Large bore variable exposure attachments can increase exposure up to 12mm in 12 increments
25 | Perforator driver for rapid and safe cranial access, designed to minimize plunge risk and ensure consistent performance across both adult and pediatric cases
26 | Fixed and rotating footed attachments for adult and pediatric use.
27 | Jacobs Chuck allows the surgeon to interchange drill bits of different sizes and types as needed during the operation.
28 | J-Latch compatible attachment for Craniotomies, Laminectomies and Transsphenoidal surgeries.
29 | Metal cutting attachment compatible enabling the safe and effective cutting of cranial fixation hardware, spinal rods, or other surgical implants
30 | Colour Coded Dissecting attachments aid in quick identification and selection of dissecting tools during surgery
31 | Unit must be antimicrobial Protection & water resistance
32 | Specify Dimensions and weight of the units
33 | Unit should be provided with all necessary accessories and consumables.
34 | Must have universal power adapter 220-240v ,50/60HZ
35 | Quality and Patient safety
35.1 | Unit must be FDA / CE certified
35.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
35.3 | Unit must comply with Electrical safety test IEC 60601-1,-2,-18 or equivalent
35.4 | Quality control ,Quality assurance and acceptance test should be done by the company, QC tools for performance verification done by End-user, if required.
35.5 | Safety assessment should be done by the company
36 | Consumables
36.1 | Vendor must include prices of all standard and optional consumables.
37 | Available Additional Options
No. | Options Description |  |  |  | REQUIRED OR NOT REQUIRED
38 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for Nurses and Physicians (End User Training) is mandatory.
3 | IEC 60601-2-38 Medical Electrical Equipment's part 2-38 Particular requirements for safety of Electrically Operated Hospital Beds
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Should have local service facility .The service provider should have the necessary equipment's recommended by the manufacturer to carry out preventive maintenance test as per guidelines provided in the service/maintenance manual
7 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
8 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
9 | Warranty : Comprehensive at least 3 Years
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Should have local service facility .The service provider should have the necessary equipment's recommended by the manufacturer to carry out preventive maintenance test as per guidelines provided in the service/maintenance manual
7 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
8 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
9 | Warranty : Comprehensive at least 3 Years