## Sheet1
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| DEFIBRILLLATOR | 25 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Offered Unit must be latest technology suitable for all types of patients and with all accessories and options
2 | Must be having multiple modes of opertaion for easy for use , to have AED, Archive Service , setup and demo ..
3 | Must be Biphasic waveform and please specify the maximum delivered energy
4 | Li ion battery with minimum 5 hrs backup / 100 shocks
5 | Please speify Charge time on a fully charged battery .
6 | Synchronize defibrillator pulse to patient's R wave and to be displayed on monitor and recorder.
7 | Paediatric paddles must be available In addition to adult Paddles .
8 | Must be able to evaluate electrode connection
9 | Unit must continously compensate for patient impedence during defibrelation.
10 | ECG/Syn/Pacing/NIBP/SPO2 with Adult and Paedetric Accessories kit
11 | ETCo2 (Capnography) must be included
12 | Unit must equipped with transcutaneous cardiac pacing
13 | Built in AED Advisory.
14 | Built-In Defibrillator tester
15 | Less than 10-12 kgs fully loaded ( includes batteries and paper rolls)
16 | Integrated Printer : Should be able to Print continuous strip of the displayed patient information and reports
17 | Offered unit can be use for both in hospital and ambulance
18 | Color monitor, >8” High Resolution Monitor
19 | Able to view monitor from different angles as well as bright light mode
20 | Able to display 4 or more waveforms simultaneously.
21 | Heart rate : Digital display 0 - 300 bpm +/- 5%
22 | User selectable Heart rate alarm
23 | Data Management : The device should capture and stores patient data, events (including waveforms and annotations), and continuous waveform and patient impedance records in internal memory. The user should be able to select and print reports, and transfer the stored information via supported communication methods
24 | Unit must be able to integrate with hospital HIS system . For that all the softwares and hardwares must be provided by vendor .
Quality and Patient safety
25 | Unit must be FDA / CE certified
26 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
27 | Unit must comply with Eletrical safety test IEC 60601-1,-2,-18 or equivalent
28 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optionals . Please specify the details clearly and all the requested spec/info must be well supported with reference documents.
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
3.1 | Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
5 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and for Nurses and Physicians (End User Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .