## Sheet1
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Portable EEG Machine | 1 | Comply | Supporting Documents Ref. | by Hospital
Offered Unit must be state of the art device and must be having latest built-in technology
| Specifications |  | Yes/No
1 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
2 | Must have montage and event setting.
3 | Must have DSA Trend graph, 3D Voltage map, Zoom window and screen comment tags.
4 | Must have at least 7 Bipolar inputs
5 | Portable on pole stand
6 | EEG amplifier must have 32 channels with built-in SpO2 and CO2 monitoring.
7 | Should have detachable Mini Junction Box for making ICU recording more effective.
8 | Must have photic stimulator for routine studies
9 | Must provide all in one PC with intuitive minimum 24” touch screen user interface .
10 | offered system must be with colour printer.
11 | The device must have capability to display up to minimum, 64 channels raw EEG waveform channels
12 | The device must be able to conduct complete minimum of 10-20 EEG attachments
13 | System capable for Long Term EEG Recording, must have on board memory features for recording upto 48 hrs.
14 | Must have a noise level not higher than 1.6µVp-p
15 | Should have Wide bandwidth to capture high frequency oscillation
16 | Sampling rate minimum of 1000hz
17 | Must perform multiple recording modes
18 | Must include on-board impedance measurement
19 | Must have customizable filters
20 | Review software-The above system must be capable of independent review stations.
21 | Unit must be provided with good Cart which is easy to move from dept to dept .
22 | Unit must be provided with PC and electrode junction box and all the needed softwares must be offered with the unit.
23 | Unit must be able to perform Quantitative analysis like aEEG, DSA and burst suppression reveals trends in EEG waveforms
24 | Software should have HL-7 connectivity.
25 | Offered unit must have Intuitive and simple user interface
26 | Powerful built-in EEG trending package
27 | Compressed view of long EEG recordings available
28 | Offered unit must have Spike and Seizure detection
29 | Designed for all types of EEG studies
30 | Ability to prune the EEG and video files before archiving for reduced storage
31 | For emergency/ night recording there should be an system headset with paste pre-filled disposable cup electrodes connected wirelessly (Blue tooth) to main EEG system.
32 | System should be able to generate Customizable report templates
33 | Remote Review option makes it possible to review live studies from a remote review station.
34 | The Remote Control option allows the user to control the recording on the acquisition station.
35 | Preferabbly offered system must be offered with video recording synchornized with EEG
Quality and Safety Standards
36 | Offered Unit must be FDA / CE marked .
37 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
38 | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
3.1 | Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and for Neuro and EEG Team (End User Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .