## C-Arm
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | C-Arm Machine | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
High frequency FDP detector intra operative surgical 3D C-Arm Machine ,DSA(Digtal subtraction angiography) with 29'' Dual ultra high resolution medical grade monitor 6 GB ,exceptional Image Quality and Efficient Maneuverability.
2 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
3 | A mobile imaging unit used primarily for fluoroscopic imaging during surgical and orthopedic procedures.Consists of a computer used to view, manipulate, store and transfer the images to fuji PACS.
4 | Integration with PACS & RIS :Unit must be able to integrate with hospital PACS system .All softwares and hardwares must be provided by vendor to do the intergration . For retrieving prior images from the PACS, the DICOM Query/Retrieve (C-FIND and C-MOVE for Patient Root and/or Study Root Model) Service Class should be support.
5 | Generator & X-ray Tube
5.1 | Microprocessor controlled high frequency digital generator
5.2 | Continuous, Intermittent & Pulsed fluoroscopy modes of operation
5.3 | Pulsed fluoroscopy ≥ 15 fps
5.4 | Auto & Manual control of fluoroscopy
5.5 | Auto High penetration for thicker patients.
5.6 | Automatic & Manual dose rate control
5.7 | X-ray tube with high anode heat capacity & cooling rate
5.8 | X-ray power ≥ 15kW
5.9 | User friendly operator console
5.1 | Integrated, Detachable, & foot switch exposure control
6 | Mobile C-Arm
6.1 | Easily maneuverable with locking facility
6.2 | wireless router
6.3 | Counter balanced with multi-directional positioning
6.4 | Maximum pivotal & orbital rotation
6.5 | Motorized height & rotational movement with safety control
6.6 | specity C-arm depth
6.7 | Specify Angulation
6.8 | specity lowest lateral postion range
6.9 | Source-Image distance ≥ 90 cm or reasonable
6.1 | Iris Collimator to limit the field to the area of interest which helps reduces radiation exposure to the patient
6.11 | Control of mA and kV parameters
7 | Detector & Image Display
7.1 | High resolution CsI digital panel detector ≥26x26 cm
7.2 | Pixel size ≤ 180µm
7.3 | Detector quantum efficiency of ≥ 70%
7.4 | Fully digital imaging chain & software
7.5 | Automatic gain & dose rate control
7.6 | Image rotation feature
7.7 | Integrated Dose rate Monitor with shortest warm-up time
7.8 | Matrix size of 2K x 2K pixel or better
7.9 | Multi Image review on dual 29” (or better) TFT flat panel monitor
7.1 | High image storage capacity
7.11 | Radiation warning light on display console
7.12 | Software such as DSA, Road mapping, edge enhancement, artifact reduction etc.
7.13 | Integrated DAP Monitor with shortest warm-up time
7.14 | Upgradeable to higher version
7.15 | PACS & DICOM interface for Image store, Query/Retrieve, Print, Work List Generation, MPPS, and archive
7.16 | Wireless DICOM compatibility
8 | Unit must be antimicrobial Protection to maintain high infection control standards
9 | Crystal-clear 2D image and high-resolution 3D reconstruction.Every tiny anatomical structures and implants become visible.
10 | Specify Weight and Dimensions of the unit
11 | Power consumption & Heat dissipation in BTU/h
12 | Quality control ,quality assurance and accaeptance test should be done by the company
13 | Safety assessment should be done by the company
14 | Software:
14.1 | Repeat /Reject Analysis software
14.2 | Statistical tool software
14.3 | Protocol analysis
15 | Accessories
15.1 | All necessary Accessories, consumables and disposables for starting-up.
15.2 | General Radiology QC tools for performance verification done by Enduser, if required.
15.3 | Lead aprons Quantity 5 ( double face 0.35 mm lead equivalen Light Weight Model
15.4 | Thyroid shield protection,Head Protection and Lead GogglesQuantity each 5 nos
15.5 | Sterile C-Arm Cover 50 Nos
16 | Trade-in (If relevant)
16.1 | Supplier should offer a Trade-in price for the following existing system:
16.2 | DESCRIPTION MANUFACTURER MODEL Sr.No YEAR OF PRUCHASE
16.3 | C-ARM X-RAY SIEMENS, GERMANY ARCADIS VARIAC 12781/647-90063888-XP9 21/07/2011
17 | Quality and Patient safety
17.1 | Unit must be FDA / CE certified
17.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
17.3 | Unit must comply with Eletrical safety test IEC 60601-1,-2,-18 or equivalent
18 | Accessories
Supplied with all standard accessories.
19 | Consumables
Supplied with all standard consumables.
20 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
19 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for OT Anesthesia technicians and Anesthesia doctors (End User Training) is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
8 | Continuous Update and Upgrading of the system and all provided software with the system including and limited to DICOM software, during the warranty (any hardware upgrade needed for such software upgrade should be included and provided with no additional cost by the vendor). It is expected from the vendor to always install the latest software versions within 4 months of its official release.
9 | Maximum time needed for parts ordering and replacement (is calculated as a system down time )
10 | Maximum response time for system faults or breakdown calls (max. time for the engineer to be at the site) to be provided.
11 | Warranty must be at laest 3 years comprehensive from date of installation and equipment acceptance.