## Sheet1
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Pulmonary Function Test | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Stae of the Art Machine must be offered with Whole body plethysmography to diagnose certain types of lung disease, such as asthma, bronchitis, ILD and emphysema. Find the cause of shortness of breath. Measure whether exposure to chemicals at work affects lung function.
2 | It should have ultrasonic flow sensor technology / latest technology
3 | Machine should be able to measure the following ;
(a) | Flow volue loop with volume timed cure. Also Static and dynamic lung Volumes and Sub divisions: FVC (ex), FEV 6, FEV 3, FEV 1, FEV 0.75, FEV 0.7, FEV 0.5, FEV1/FVC, FEV1/IVC, FEV1/VC, FEV0,5/FVC, FEV0,5/IVC, FEV1*30, PEF, MEF 25, MEF 50, MEF 75, MEF 25-MEF75, AEX, t ex, EV, FVC (in), FIV1, FIV1/VC, PIF, MIF 25, MIF 50, MIF 75, MEF50/MIF50
(b) | Pre and post bronchodilator with comparision,percentage variability
(c) | Single breath DLCo and KCO
(d) | Single breath helium dilution FRC lung volumes
4 | Must be having impulse Oscillometry
5 | Machine Should be able to measure airway resistence
6 | Parameters should be measured with highest accuracy reproducibility
7 | The System should have temperature , barometric and humidity sensors
8 | All Sensors should preferably be non-consumable or their life should be specified in the offer.
9 | Real time flow volume and volume time curves (both inspiratory and expiratory) on computer screen.
10 | Actualized fletcher curve on report
11 | Country Specific predicted values Selectable
12 | Ethinic correction factors
13 | Children incentive available
14 | Gas diffusion should have the following analyzers ; CO diffusion TLco, Kco, TLC, RV, FRC, VA , t diff, Hb, VIN (ch4)
15 | Helium technique (Ultrasonic MOI mass/Thermal conductivity determination. Nitrogen washedout method
16 | CO ( Should be non-dispensing Infrared(NDIR) or latest
17 | O2 (With Solid State electrolyte technology / latest technology .
18 | System Should have portable with in-built touch Screen monitor
19 | System Should have LAN port for data sharing and networking
20 | It Should have true infection control and low maintenance cost
21 | Compact with single cylinder for all the test
22 | It Should have mobile cart
23 | Machine should be able to do provocation test
24 | For aged patient it would be preferable , if the machine would have Wheelchair access option.
25 | Should meet ATS/ERS/(2005-standards) guidelines ; It should provide necessary accessories and consumables.
26 | Respiratory mechanics , ( MIP/MEP) and airway resistence.
27 | The system must be supplied with Diffusion test gas cylinder along with regulator , bacterial filter Silicon mouthpieces Nose Clips .
28 | It should provide necessary accessories and consumables.
SYSTEM AND APPLICATION SOFTWARE
29 | Should be supplied with the equipment . The management software Should be ;
30 | Windows 10 or latest Environments
31 | Multifunction laser printer with wireless connectivity
32 | Machine should be able to integrate with hopistal HIS system
33 | Compatible software to upgrade for cardiopulmonary exercise test
Quality and Patient safety ( Support documents must be submitted)
34 | Unit must be Quaity certified and safe for patient use as per internationals standards ,
35 | Unit must meet ATS/ERS standards.
36 | Unit must be FDA / CE certified
37 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
38 | Unit must comply with Eletrical safety test IEC 60601-1,-2,-18.
OPTIONS
It's preferable that unit will be provided with option that doctors can view the graphs and patient readings in their clinic's computer for further analysis.
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
3.1 | Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and End user / application training must be done for Nurses and for Physicians .
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .