## OT Table
Indicative Specifications / Evaluation Form
Item No. | Operating Table for the OT | Qty. | To Be Filled By Vendor | To Be Filled By Hospital
|  | 2 | Supporting Documents Reference | Met
1 | Technical Specs |  |  | Yes | No
1.1 | A mobile operating table for general surgery , robust and rust proof and impact resistence.
1.2 | Frame/design
1.2.1 | The system should be easy to use and adjustable for heights, trendelenburg, lateral tilt.
1.2.2 | It should be opertable on mains and battery.
1.2.3 | It should be equipped with four lockable, double swivle castors for easy manoeuvering.
1.2.4 | Base cover should be made of material that is resistant to scratches, impact, fractures and disinfectants.
1.2.5 | Supporting bars should all be made of stainless steel.
1.2.6 | Table top should not have any crossbars allowing for C-arm access.
1.2.7 | Should have stepless electrical adjustment for all movements and positions: Height, Back section, Leg Section, Trendelenburg , reverse-Trendelenburg, etc.
1.2.8 | It should have SFC padding of at least 70-80 mm.
1.2.9 | The height of the operating table should be adjustable from a minimum of 600 mm to 1000 mm.
1.2.10 | It should facilitate an easy and quick way for module replacements.
1.2.11 | Weight capacity should not be less than 450 Kg.
1.2.12 | All functions of the operating table should be controllable with a key pad and also through hand control.
1.2.13 | It should have radiotransluscent top for carrying out X-Ray at the bedside and facilitates an easy insertion of the X-Ray cassette.
1.2.15 | Operating table to include additional head rest, upper back plate, lower back plate, extension plate, seat plate and leg plates.
1.2.16 | Length of table without accessories should not be less than 800 mm.
1.2.17 | Width of seat should not be less that 500 mm.
1.3 | Adjustments
1.3.1 | Lateral tilt (left/right) approx. 20°
1.3.2 | Lower back plate arrox. +80°/-40°
1.3.3 | Leg plate approx. +80°/-90°
1.3.4 | Trendelenburg, reverse trendelenburg approx. +30°/-30°
1.3.5 | Head rest approx. +40°/-40°
2 | Accessories
2.1 | Vendor to quote supply all accessories that will required for full intended functioning of the equipment
2.2 | Mandatory accessories to be included ;
2.3 | Arm board -Pair
2.4 | Extension Segment
2.5 | Head rest
2.6 | Leg Plates - Pair
2.7 | Ansthesia Screen
2.8 | Patient Body Strap
3 | Consumables
3.1 | Vendor must include prices of all standard and optional consumables.
4 | Available Additional Options
No. | Options Description |  |  | REQUIRED OR NOT REQUIRED
Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for Nurses and Physicians (End User Training) is mandatory.
3 | IEC 60601-2-38 Medical Electrical Equipments part 2-38 Particular requirements for safety of Electrically Operated Hospital Beds
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Should have local service facility .The service provider should have the necessary equipments recommended by the manufacturer to carry out preventive maintenance test as per guidelines provided in the service/maintenance manual
7 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
8 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
9 | Warrenty : Comprehensive atleast 3 Years