## OT Table
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To Be Filled
| Operating Table for the OT | 2 | Comply | Supporting Documents Reference | By Hospital
| Specifications |  | Yes/No (If Yes Specify)
The state -of -the-Art Operating Table for the OT ,advance latest model ,The operation theater table should support patients during all Surgical procedures and not limited to Table Design, Size and Adjustability,Mobility and Stability,Positioning and Trendelenburg,Safety Features and Accessories and Attachments Etc.
1 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
1.1 | A mobile operating table for general surgery , robust and rust proof and impact resistence.
1.2 | Frame/design
1.2.1 | The system should be easy to use and adjustable for heights, trendelenburg, lateral tilt.
1.2.2 | It should be opertable on mains and battery.
1.2.3 | It should be equipped with four lockable, double swivle castors for easy manoeuvering.
1.2.4 | Base cover should be made of material that is resistant to scratches, impact, fractures and disinfectants.
1.2.5 | Supporting bars should all be made of stainless steel.
1.2.6 | Table top should not have any crossbars allowing for C-arm access.
1.2.7 | Should have stepless electrical adjustment for all movements and positions: Height, Back section, Leg Section, Trendelenburg , reverse-Trendelenburg, etc.
1.2.8 | It should have SFC padding of at least 70-80 mm.
1.2.9 | The height of the operating table should be adjustable from a minimum of 600 mm to 1000 mm.
1.2.10 | It should facilitate an easy and quick way for module replacements.
1.2.11 | Weight capacity should not be less than 450 Kg.
1.2.12 | All functions of the operating table should be controllable with a key pad and also through hand control.
1.2.13 | It should have radiotransluscent top for carrying out X-Ray at the bedside and facilitates an easy insertion of the X-Ray cassette.
1.2.14 | Operating table to include additional head rest, upper back plate, lower back plate, extension plate, seat plate and leg plates.
1.2.15 | Length of table without accessories should not be less than 800 mm.
1.2.16 | Width of seat should not be less that 500 mm.
1.2.17 | Unit must be antimicrobial Protection
1.2.18 | Specify Dimensions and weight of the units
1.2.19 | Unit must have universal power adapter 220-240v ,50/60HZ
1.2.20 | Specify the power consumption for the offred units and heat dissipation(BTU/hr)
1.2.21 | Unit must be available for trail/Demo use
1.3 | Adjustments
1.3.1 | Lateral tilt (left/right) approx. 20°
1.3.2 | Lower back plate arrox. +80°/-40°
1.3.3 | Leg plate approx. +80°/-90°
1.3.4 | Trendelenburg, reverse trendelenburg approx. +30°/-30°
1.3.5 | Head rest approx. +40°/-40°
2 | Accessories
2.1 | Vendor to quote supply all accessories that will required for full intended functioning of the equipment
2.2 | Mandatory accessories to be included ;
2.3 | Arm board -Pair
2.4 | Extension Segment
2.5 | Head rest
2.6 | Leg Plates - Pair
2.7 | Anesthesia Screen
2.8 | Patient Body Strap
3 | Trade-in (If relevant)
3.1 | Supplier should offer a Trade-in price for the following existing system:
3.2 | DESCRIPTION MANUFACTURER MODEL YEAR OF PRUCHASE QUANTITIES
3.3 | SURGICAL TABLE MAQUET, GERMANY ALPHAMAXX 1133.12B1 07/09/2011 02
4 | Quality and Patient safety
4.1 | Unit must be FDA / CE certified
4.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
4.3 | Unit must comply with Electrical safety test IEC 60601-1,-2,-18 or equivalent
4.4 | Quality control ,Quality assurance and acceptance test should be done by the company, QC tools for performance verification done by End-user, if required.
4.5 | Safety assessment should be done by the company
5 | Consumables
5.1 | Vendor must include prices of all standard and optional consumables.
6 | Available Additional Options
No. | Options Description |  |  |  | REQUIRED OR NOT REQUIRED
7 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for Nurses and Physicians (End User Training) is mandatory.
3 | IEC 60601-2-38 Medical Electrical Equipments part 2-38 Particular requirements for safety of Electrically Operated Hospital Beds
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Should have local service facility .The service provider should have the necessary equipments recommended by the manufacturer to carry out preventive maintenance test as per guidelines provided in the service/maintenance manual
7 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
8 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
9 | Warrenty : Comprehensive atleast 5 Years