## Sheet1
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| ENDOSCOPY WASHER SYSTEM | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
Please specify the below details ;
1 | Number of Basin Available ( It is mandatory that unit must have two independent , asynchronous reprocessing basins so in case if one unit is defective the other can work independently .
2 | Cycle Time (it should be least as possible)
3 | Num. Cycle Per Hour
4 | Water and Disinfectant temperature Range
5 | Disinfection mode ( Programmable, Self and automated disinfection)
6 | Availbility of required Chemicals , chemicals must be approved for safe use and Please mention the compatable /types of disinfectant for the offered unit model.
7 | Continous Leak test
8 | Self-disinfects
9 | Avaibility of thermal Printer
10 | Vapor management system
11 | Test Strips Compatability - To test high-level disinfectant concentration with test strips.
12 | Automated leak testing for endoscopes must be available
13 | Scopes Compatability : It is important that offered system must be easy for use and compatable with mulitple scopes and brands.
14 | All the relevant connectors must be availble and provided to attach different endoscopes .
15 | Please specify ; Input Water and Electrical Requirements : (a) Input Flow Rate (b) Water temperature (c) Water hardness (d) Pre-filters (e) Electrical Voltage and Current Ratings
Quality and Patient safety
16 | Unit must be FDA / CE certified
17 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
18 | Unit must comply with Eletrical safety test IEC 60601-1,-2,-18 or equivalent
19 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optionals . Please specify the details clearly and all the requested spec/info must be well supported with reference documents.
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
3.1 | Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and for Nurses and Physicians (End User Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .