## Sheet2
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| HIGH END ICU BED | 3 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Bed should be FDA/CE Approved
2 | Bed should be made at good quality standards in terms of Efficiency , Quality , safety and Connectivity .
3 | Electrical Bed with fully integrated Air mattress designed for intensive Care Areas, with Battery Backup
4 | Connectivity ready with NurseCall
5 | One button function for Bed up/down, backrest up/down, knee up/down, Trendelenburg/revs.Trend
6 | In Bed Scale weighing in Kg
7 | 30°/45° Head Angle Alarm
8 | Retractable Foot Control
9 | Full Chair Position with possibility of chair Egress
10 | Emergency CPR from both sides
11 | Bed Exit Alarm with different modes
12 | Provides rotation therapy within the Mattress for patients at risk of pulmonary complications.
13 | Have percussion and vibration therapies in the mattress to help clearing secretions from the lungs
14 | Lateral patient Turn function
15 | Central brake system and steering mechanism located on either side of the bed with clearly marked pedals. With Brake not set Alarm
16 | Should have Auto Contour Feature which allows the bed to prevent the patient from sliding down in the bed.
17 | Caregiver Siderail Controls on both sides of the bed
18 | All Frame Functions Electrical
19 | Split Side Rails, with easy Folding down and One Hand Release Siderails
20 | Radiolucent Sleep Deck, with X-Ray sleeve under patient’s chest, and C-Arm compatibility
21 | Safe Working Load minimum 250 Kgs
22 | Color Touch Screen Interface with digital angle indicator
23 | Roller Bumper System
24 | Angle Indicator On both sides
25 | Night Light
26 | 4 Sockets for I.V pole
27 | Bed not in Lowest position indicator
28 | Microclimate technology in the mattress
29 | Head & Foot Board
30 | Drainage Bag Holders
31 | Line Manager
32 | Transport Handles
33 | 2 Vertical Oxygen Tank Holders
33 | All positions must be lockable with secure key
(a) | Articulation specification approx.
(b) | Head 0°-67°
© | Thigh 0°-30°
(d) | Foot 0°-70°
( e) | Trendelenburg 13°
(f) | Reverse Trendelenburg -13°
(g) | I.V pole
34 | Integrated pattient mattress is required ( Powered )
Which address immobility associated complications. (b) Help Protect Skin integrity ( c) Promote patient Comfort .
35 | The unit must comply with all applicable ISO/CE/MDD 93/42 or equivalent.
36 | The unit must comply with quality standard including ISO 13485 or quivalent .
37 | The Unit must be hospital grade , that meets EN 60601-1 performance and safety requirements
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and for Nurses and Physicians (End User Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .