## Sheet1
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
1 | CRRT Machine (Continuous renal replacement therapy) | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Easy to handle and maintain
2 | Microprocessor/microcontroller controlled user interactive menu with operating and malfunction removal instructions on the display screen
3 | The system should be compatible with Hemodialysis/Haemofiltration
4 | Minimum Four pumps, one each for Blood, Dialysate, Replacement fluid and able to perform SCUFF, CVVH, CVVHD, CVVHDF &PLASMAEXCHANGE
in addition to the above mentioned modes of operation the offered system must be able to perform ;
(a) | SLEDD (Sustained low efficiency dialysis)
(b) | Ci-Ca therpy ( Citrate Anticoagulation therapy) or any similar mode
© | SEPSIS Filteration
(d) | MARS (Molecular Adsorbent Recirculating system) - (OPTIONAL)
€ | CO2 Removal
5 | Clear touch screen TFT/LCD Monitor
6 | Blood pump speed of 10-450 ml/min.
7 | Close blood circuit to prevent air to blood interface
8 | Short preparation and priming program and ready to start treatment less then 30 minutes.
9 | Arterial pressure range: Preferably (-) 250 mmHg +/- 50mmHg.
10 | Venous pressure range Preferably (-) 50 mmHg to 350mmHg.
11 | Pre Fitter Pressure Preferably (-) 50mmHg to +450mmHg.
12 | Effluent Pressure Preferably 350mmHg +/- 50mmHg
13 | Programmable Substitution solution flow rate Approx : 0-5000 mL/Hr
14 | Dialysate flow rate:0-2500 mL/Hr
15 | Programmable Effluent including CVVHDF Flow Rate Preferably : 60-10000 mL/Hr
16 | Integrated heparin pump with a flow rate of 0. 0.5 ml-5 ml/Hr. Bolus facility range 0.5mL. Bolus frequency immediate 1-24 hrs.
17 | Capable of changing therapies
18 | Four weighing scales to control system with gravimetric fluid mechanism and weighing capacity of at least 0-11kg, 7g deviation for a solution bag of 5200g (equivalent to 0.14%) or please specify if there is any equiivalent control mecahnism availble other then weight measurement.
19 | Fluid/Blood warmer for blood/dialysate warming temp range app 33-40 degC(+/- 3 deg C)
20 | Ultrasonic or latest technology air bubble detector
21 | Alarm in case of blood leak, air in line pressure limit violation, empty dialysate / Replacement bag, full effluent bag, and advisory alarms in case of excessiveTMP and filter clotting.
22 | Output for Printer, PC connectivity, and Data acquisition should be avaible.
23 | Unit must be able to integrate with hospital HIS system . For that all the softwares and hardwares must be provided by vendor .
Quality and Patient safety
24 | Manufacturer/Supplier should have ISO certification for quality standards.
25 | Should be FDA/CE/ UL or BIS approved product
26 | Shall comply with IEC 60601-2-16 SAFETY requirement of medical electrical equipment part 2- particular requirements for the safety of Haemodialysis equipment.
27 | Technical offer must be clear in terms of standards , which will be offered with the machine and Optionals . Please specify the details clearly and all the requested spec/info must be well supported with reference documents.
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for ICU Physicians , Nurses , technicians and technical training for biomed dept is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .