## Sheet1
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| PORTABLE ECHO ULTRASOUND (FOR OPD Cardiology ) | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
1 | Portable ultrasound Unit
2 | Monitor Screen minimum 15 inches
3 | Durable Trolly with three probe connectors
4 | Having B/M color and PW/CW
5 | Must be supplied with HDMI /DVI adapter
6 | Frame rate/Sampling rate must not be less then 150 frames/sec
7 | Must be supplied with 500GB HDD or more
8 | Unit must be be able to provide at least 2-3 hrs of battery back-up
9 | Speckle suppresion
10 | Spatial Compounding
11 | Anatomical M-Mode
12 | Full DICOM and HL7 compatability
13 | Phased array adult cardiac probe 1 to 5 Mhz
14 | ECG leads and Color Printer
15 | Must be supplied with full cardiac package
16 | Linear Probe 3Mhz to 9Mhz
18 | Paedeatric Sector Probe 3Mhz to 8Mhz
19 | Meonatal Sector Probe 4Mhz to 12Mhz
20 | Speckle tracking
21 | Unit must be able to integrate with Hospital HIS system
Quality and Patient Safety Compliance : ( Support documents must be supplied )
22 | Unit must be FDA / CE marked .
23 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
24 | System must comply with ISO ; IEC Medical equipment Electrical Safety IEC 60601-1..
2 | Accessories
Supplied with all standard accessories.
3 | Consumables
Supplied with all standard consumables.
4 | Avialable Additional Option
No. | OPTIONS DESCRIPTION
Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and for Nurses and Physicians (End User/Application Training) is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed.Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warrenty : Comprehensive minimum 3 Years warrenty .