## Oral Liquid Unit-Dose Packaging
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Oral Liquid Unit-Dose Packaging Machine | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the Oral Liquid Unit-Dose Packaging Machines and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
2 | The system shall be a fully automated or semi-automated oral liquid unit-dose packaging machine designed for hospital pharmacies, long-term care facilities, and pharmaceutical repackaging applications.
3 | Please specify make, model and origin of the machine.
5 | The machine shall package bulk oral liquid medications into individually sealed, barcoded unit-dose cups or containers.
6 | High-barrier packaging material providing protection against moisture and light.
7 | The machine shall be supplied complete with all necessary hardware, software, accessories, and consumables required for operation
8 | Packaging Performance
8.1 | Production capacity: Minimum 20 unit doses per minute.
8.2 | Filling accuracy: ±1% or better.
8.3 | Capable of packaging oral liquid medications of varying viscosities.
8.4 | Automatic filling, sealing, labelling, and cutting process.
8.5 | Ability to package different batch sizes ranging from small validation runs to high-volume production
8.6 | Automated calculation of Beyond-Use Dates based on default configuration by end-user
9 | Packaging Format
9.1 | Unit-dose cups available in multiple sizes, including:7 mL,15 mL,25 mL & 35 mL
9.2 | Tool-free changeover between cup sizes.
9.3 | High-barrier packaging material providing protection against moisture and light.
9.4 | Tamper-evident sealing mechanism.
9.5 | Packaging materials shall comply with USP Class A requirements or equivalent.
9.6 | Maintain a complete, unalterable, and time-stamped electronic audit trail to track all user activities, including but not limited to logins/logouts, system calibrations, and any modifications made to label templates.
10 | Dosing System
10.1 | Integrated precision syringe pump or equivalent dosing technology.
10.2 | Use of FDA-grade disposable tubing and dosing accessories.
10.3 | Automatic low-liquid detection and alarm system.
10.4 | Adjustable filling volume and pumping speed through software interface.
10.5 | Safety Sensor: "No Cup – No Fill" optical logic to prevent pump dispensing if a cup is missing.
11 | Labelling and Barcoding
11.1 | Integrated thermal printer for direct package labelling.
11.2 | Support for: Linear barcodes,2D barcodes,GS1 barcodes,
11.3 | Printing of: Drug name,Strength,Lot number, Expiry date, Packaging date, Barcode information, Optional barcode verification system. Support Tall-Man lettering for look-alike/sound-alike medications.
11.4 | Support Tall-Man lettering for look-alike/sound-alike medications.
11.5 | GS1 QR Codes Support for multi-field data embedding in a single scan.
12 | Windows-based packaging and labelling software.
13 | Drug database management capability.
14 | User access control with password protection.
15 | Audit trail and reporting functionality.
16 | Master Drug Data Syncing from Trak Care Intersystem(HIS)formulary
17 | Network connectivity for pharmacy information systems.
18 | Capability to store product-specific packaging parameters.
19 | Exportable reports in standard formats
20 | Colour touchscreen display (minimum 7 inches) or better .
21 | Multilingual user interface.
22 | unit must be available for demonstration
23 | Lightweight and compact design suitable for clinical rooms.
24 | Easy to clean surfaces and disinfect able parts.
25 | Unit must be antimicrobial protection
26 | Power supply: 220–240V, 50/60 Hz
27 | Please specify the weight and unit must be provided with good quality castors .
28 | The Machine shall support seamless integration with the Hospital Information System (HIS) to enable automated synchronization of the drug formulary and ensure accurate dispensing operations.Below are the technical Integration Requirements
28.1 | The proposed solution shall support standard HL7 Version 2.5 messaging for integration with the Hospital Information System (HIS).
28.2 | The machine shall support automated Drug Formulary synchronization using the HL7 MFN message.
28.3 | The integration shall allow the machine to receive new drug records as well as updates, modifications, activations, and deactivations of formulary items from the HIS.
28.4 | The supplier shall be fully responsible for designing, implementing, testing, and commissioning the HL7 integration with the Hospital Information System.
28.5 | All software, interface components, middleware, interface engines, licenses, and configuration required for successful integration shall be included within the supplier's proposal.
28.6 | The supplier shall provide on-site technical support during installation, testing, Go-Live, and stabilization until successful completion of the integration project.
28.7 | The supplier shall provide complete technical documentation.
28.8 | The supplier shall ensure that the integration complies with international HL7 standards and healthcare interoperability best practices and shall work collaboratively with the Hospital Information System vendor to achieve successful end-to-end integration.
28.9 | The supplier shall provide all software licenses, interface licenses, middleware licenses, communication licenses, APIs, drivers, connectors, and any other software components required for successful integration with the Hospital Information System (HIS).
28.3 | All licenses required for the development, testing, production, and future operation of the integration shall be included in the supplier's proposal without any additional cost to the hospital unless explicitly stated in the commercial offer.
29 | Quality and Patient safety
29.1 | Device should be CE and FDA approved
29.2 | Offered System must comply with Electrical safety IEC 6060-1
29.3 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
29.4 | Electrical leakage and overload protection
30 | Accessories
30.1 | All necessary accessories
31 | Consumables
31.1 | Supplied with all standard consumables-Disposable
32 | Available Additional Option
No. | OPTIONS DESCRIPTION
33 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Pharmacy doctors and other associated staff and Service training for Biomedical Engineering Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warranty .