## Air pendrive drill system
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To Be Filled
| Orthopedic drill system | 1 | Comply | Supporting Documents Reference | By Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the Orthopedic drill system and Optional items . Please specify the details clearly and all the requested specification info must be well supported with reference documents.
2 | Specify Origin, Make and Model of the unit
3 | Offered model should not have a recall, and in case the offered model is purchased and there is a recall during warranty then the supplier has to provide a replacement within 5 days with the same features of the recalled model.
4 | The offered models of each part of the complete system should be from same manufacturer brand and Should have an authorized distributor in UAE.
5 | The Required system should have the below specifications.
5.1 | Battery-driven orthopedic drill system for Trauma, arthroplasty, spine, CMF, sports medicine.
5.2 | Dual trigger modular handpiece compatible with full range of attachments covering all orthopedic indications
5.3 | Operate continuously during major orthopedic procedures.
5.4 | Permit rapid attachment exchange during surgery.
5.5 | Maintain torque under heavy drilling and reaming load.
5.6 | Allow precise control at low and high speeds.
5.7 | Minimize thermal bone necrosis through efficient cutting performance.
5.8 | Should be suitable for repeated sterilization cycles without degradation.
5.9 | Maximum speed not less than 1400RPM, maximum torque not less than 12NM
5.1 | Oscillation Speed Minimum 11,000 osc/min
5.11 | Modular Drive Speed > 20,000 rpm without attachment
5.12 | Variable speed trigger control with forward and reverse rotation.
5.13 | Brushless sterilizable motor
5.14 | Battery capacity not less than 30Wh enough to operate complex challenging long cases
5.15 | Battery charge time ≤ 60 minutes
5.16 | specify the battery backup in ____mins
5.17 | Overheat and overload Safety protection
5.18 | Low-noise operation preferred
5.19 | Low-noise operation preferred
5.2 | Ergonomic lightweight design to minimize surgeon fatigue and improve handling.
6 | Attachment & Functional Accessories:
6.1 | Power tool Handpiece(Minimum 0–1450 rpm with drill attachment) | 1
6.2 | Power Battery(Lithium-Ion rechargeable)-Integrated battery indicator status display | 2
6.3 | Battery Charger | 1
6.4 | Reciprocating Saw Attachment | 1
6.5 | Sagittal Saw Attachment | 1
6.6 | Quick Coupling for DHS/DCS Triple Reamers | 1
6.7 | Quick Coupling for Kirschner Wires | 1
6.8 | Attachment for Acetabular and Medullary Reaming | 1
6.9 | Drill Chuck (Drilling Speed) | 1
6.1 | AO/ASIF Quick Coupling | 1
6.11 | Trinkle Quick Coupling (Reaming Speed) | 1
6.12 | Trinkle Quick Coupling (Drilling Speed) | 1
6.13 | Hudson Drill Attachment | 1
6.14 | Hudson Ream Attachment | 1
6.15 | Instruments tray with lid | 1
7 | Steam autoclavable at 134–135°C
8 | Unit must be antimicrobial Protection and water resistance(IPX4 or equivalent)
9 | Specify Dimensions and weight of the units
10 | Unit should be provided with all necessary accessories and consumables.
11 | Must have universal power adapter 220-240v ,50/60HZ
12 | Quality and Patient safety
12.1 | Unit must be FDA / CE certified
12.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
12.3 | Unit must comply with Electrical safety test IEC 60601-1,-2,-18 or equivalent
12.4 | Quality control ,Quality assurance and acceptance test should be done by the company, QC tools for performance verification done by End-user, if required.
12.5 | Safety assessment should be done by the company
13 | Consumables
13.1 | Vendor must include prices of all standard and optional consumables.
Specific consumables for different length attachment, offering tailored performance and optimized efficiency.
14 | Available Additional Options
No. | Options Description
15 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and for Nurses and Physicians (End User Training) is mandatory.
3 | IEC 60601-2-38 Medical Electrical Equipment's part 2-38 Particular requirements for safety of Electrically Operated Hospital Beds
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 10 years of operation and maintenance of the equipment supplied.
6 | Should have local service facility .The service provider should have the necessary equipment's recommended by the manufacturer to carry out preventive maintenance test as per guidelines provided in the service/maintenance manual
7 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
8 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
9 | Warranty : Comprehensive at least 3 Years