## MRI Machine
Indicative Specifications / Evaluation Form
Item No. | Item Description | To Be Filled By Vendor |  | To be filled
| MRI Machine 3 Tesla | Comply | Supporting Documents Ref. | by Hospital
| Specifications | Yes/No (If Yes Specify)
State of Art High Field whole body MRI system Superconductive Magnet (3 Tesla) with wide bore 70-75 cm size,Maximum available channels is preferred with superconducting 5-7G Integrated latest AI technology ,Advanced Imaginh Applications,low noise, fast imaging, Silient scan with enhanced patient experience, having all the anatomy specific coils and software for all types of clinical software application including and not limited to Neurology, Paediatrics, Oncology, Musculoskeletal, Abdomen, liver, Body, Prostate, Vascular, All MR angiography,Functional studies,cardiac .. etc. all features expected from each of these application to be offered as standard, other avialble applications should be clearly mentioned as optional.
Installation and B₀ main field
Technical offer must be clear in terms of standards , which will be offered with the MRI machine and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
Please specify make, model, and origin of the machine.
Faraday cage(shild) included, Technical room, UPS, MRI chiller.
All room preparation and system installation requirements (to be clarified and agreed on after vendor initial site visit)- full room layout required.
Installation should be done by original manufacturer 3 T trained staff
1 | All pre-installation requirements should be submitted along with the quotation including suggested room dimensions etc.
1.1 | Power Supply Capacity and Specifications
a | Minimum Space Requirement for your system
b | Minimum installed area (m2)
c | Faraday cage(shild) included, Technical room, UPS, chillers
d | All room preparation and system installation requirements (to be clarified and agreed on after vendor initial site visit)- full room layout required.
e | Installation should be done by original manufacturer 3 T trained staff
1.2 | Scanner FDA nod
1.3 | Scanner CE certification
1.4 | Field strength:
1.5 | Dimensions of magnet bore
a | Patient Aperture Width cm at narrowest: (width x height)
b | Magnet bore length (w/o covers):
1.6 | Fringe field
a | Is the magnet equipped with magnetic shielding?
b | Distance of the 0.5-mT line from isocenter
c | horizontal axis (x):
d | vertical axis (y):
e | longitudinal axis (z):
1.7 | Shielding against external interferences
a | Is the magnet shielded against external interferences? (e.g. moving ferromagnetic objects, transformers)
b | Is shielding against external interferences active during a scan?
c | Description:
1.8 | Magnet coils
a | Number of magnet coils:
b | Are magnet coils actively shielded?
1.9 | Main magnetic field Homogeneity of <1 ppm over a 40 cm DSV
a | Max Field of view > 50 cm
b | Inside a volume of 55 x 55 x 50 cm³ or equilvalent
1.10 | Long term stability of the magnetic field:
1.11 | Helium Consumption
a | Zero helium boil-off technology
b | Helium refill yearly
c | Tank capacity of liquid helium at 100% level:
1.12 | Shim System for the Installation
a | Passive Shim for improvement of the field homogeneity (e.g. with iron):
b | Active Shim for improvement of the field homogeneity(Active shimming preferred)
c | Number of independent shim terms with the active shim (e.g. 8, for the 3 linear terms and for the 5 nonlinear terms):
1.13 | Patient-specific Shim
a | Is active shimming for a patient-specific improvement of the field homogeneity possible?
b | Number of independent shim channels with the active shim (e.g. 3 for the 3 linear channels):
c | Fully automated patient-specific shimming:
d | Typical duration of the automatic shim:
1.14 | Advanced shim system for patient-specific shim
a | Active shim system (with non-linear shim terms) for a patient-specific improvement of the field homogeneity:
b | Number of independent shim channels with the advanced active shim system (e.g. 3 for the 3 linear channels):
c | Fully automated patient-specific shimming:
d | Typical duration of the automatic shim:
1.15 | RF CAGE & STANDARD ACCESSORIES
a | RF cage - RF room with one Door and one View window (min 140 x 100) - state full technical specifications of the offered RF cage.
b | The maximum acoustic noise level expected outside the MRI room (state the technical specification of the door of MRI room offered - with its sound (acoustic noise) proof
c | MR Room (zone 4) and zone 3 Oxygen Monitors
d | Door Mounted Metal Detector
e | Safety marking plate with Alarming System in case of Chiller/ Magnet/Power down.
f | 2 Quench buttons (in zone 4 and above MR console)
g | Measuring and Adjustment Kit : Includes valves, Flow meter, water filter, shunt valve & fittings
1.16 | In Bore patient comfort, and monitoring features
a | Feet-First or head -first entry
b | In-bore lighting system
c | Audiovisual device for patient comfort and funtional research
d | In-bore patient ventilation system
e | Patient accessories: Dedicated head arm support, Table mattress set, Head/leg support, Knee support, Positioning wedges, Small foam wedges and pads, Set of sandbags, positioning pads, Immobilization straps, Set of patient fixation straps etc
f | Self-storing non-ferrous IV pole
g | Comprehensive set of pediatric accessories : Includes Baby support mattress for small infants, Anterior coil frame etc.
2. Gradient System
2.1 | The required gradient performance has a peak gradient field strength of at least 45 mT/m per axis and a Maximum slew rate of ≥ 200 T/m/s per axis simultaneously or Equivalent (true physical value)
For gradient systems with several operational modes, the specifications for each gradient mode have to be provided.
2.1.1 | Describe the shape of the gradient linearity volume
2.1.2 | Are the gradient coils actively shielded?
2.1.3 | Maximum gradient amplitude (per axis): ≥ 60 mT/m (true physical value)
2.1.4 | Maximum slew rate
2.1.5 | Minimum rise time from 0 to max. gradient amplitude:
2.1.6 | Duty cycle:
2.1.7 | Manufacturer of the gradient coil: Values must be true physical per-axis values (not vector/effective).;
2.1.8 | Maximum output current per axis of the Gradient Power Amplifier (GPA) horizontal axis (x, "left-right"), vertical axis (y, "anterior-posterior"), longitudinal axis (z, "head-feet")
2.1.9 | Maximum output voltage per axis of the Gradient Power Amplifier
horizontal axis (x, "left-right"), vertical axis (y, "anterior-posterior"), longitudinal axis (z, "head-feet")
2.1.10 | Maximum output voltage and output current available in one single gradient pulse
2.1.11 | Manufacturer and type of the gradient amplifier:
2.1.12 | Maximum Field-of-View for the above-mentioned gradient specifications (depending on application) horizontal axis (x, "left-right"), vertical axis (y, "anterior-posterior"), longitudinal axis (z, "head-feet")
2.1.13 | Minimum TR ms Spin echo
2.1.14 | Minimum TR ms 2D gradient echo
2.1.15 | Minimum TR ms 3D gradient echo
3 | Radio Frequency (RF) Transmitter System B1 field
3.1 | Signal generation and processing, RF technology?/Fully digital RF architecture.
3.2 | Frequency resolution of the RF synthesizer:
3.3 | Bandwidth of the RF transmitter:
3.4 | Max. power of the transmitter amplifier:
3.5 | RF Multi Transmit Capability
3.6 | Is the amplitude and phase transmission optimized for dedicated body regions
3.7 | Upgradeable to higher number of receive channels without major system replacement.
3.8 | Is the transmission system integrated in the magnet housing?
4. RF Receiver System
4.1 | Are the receive components integrated into the magnet housing?
4.2 | Maximum number of channels. Minimum of 32 channels, anything higher is preferred
Signal receiving and processing
4.3 | Number of independent receiver channels that can be used simultaneously in one single scan and in one single FOV, each generating an independent partial image
4.4 | Partial Images to proof independent channels to be provided for at least 5 body regions (e.g. Head, Head/neck, Body, ...)
Max. receiver bandwidth of each receiver channel:
4.5 | Do all channels have the same bandwidth?
4.6 | Sampling rate of each ADC:
Receiver signal resolution:
4.7 | Noise of the preamplifier:
4.8 | Max. number of Coil Elements
Max,number of available coils that can be used in FOV
4.1 | Max. number of simultaneously connected coil elements:
4.9 | Can all of these coil elements be seamlessly used in one study without coil changes and without repositioning of the patient?
Provide an example of the max. number of coil elements that can simultaneously be connected and used in one examination without coil or patient repositioning:
5 | RF Coils
5.1 | Special volume and surface coils for the following applications shall be available with the system. The required coils have to be best in the market with the maximum number of elements and to be included in the quote.
5.1 | Applications
5.1.1 | Multi-channel array coils for high-speed Parallel Acquisition Techniques:
5.1.2 | Head
5.1.3 | Head-Neck combined, Neuro-Vascular
5.1.4 | Cervical Spine, Thoracic Spine, Lumbar Spine, Whole Spine
5.1.5 | Whole CNS (Head + Whole Spine):
5.1.6 | Chest, Heart:
5.1.7 | Abdomen:
5.1.8 | Pelvis:
5.1.9 | Whole Torso:
5.1.10 | Hip:
5.1.11 | Knee, Ankle and Foot:
5.1.12 | Shoulder:
5.1.13 | Elbow:
5.1.14 | Wrist:
5.1.15 | Breast:
5.1.16 | Lower Extrimity coils for peripheral with PAT compatability Angiography procedure
5.1.17 | Whole-Body scanning with high-SNR surface coils:
5.1.18 | Which high-SNR surface coils and/or coil combinations can be used simultaneously for whole-body scanning?
5.2 | Coil Technology
5.2.1 | Is it possible to connect multiple coils simultaneously?
5.2.2 | How many different coils can be connected and used simultaneously?
5.2.3 | Is it possible to combine coil elements from different coils, i.e. use coil elements from different coils simultaneously in one single scan?
5.2.4 | May certain coils stay on the patient table for the next examinations, in order to minimize coil changes?
5.2.5 | Please list these coils and name possible limitations:
5.2.6 | Are connected surface coils automatically detected by the system?
5.2.7 | Does this apply to all surface coils?
5.2.8 | Is the position of the coils in relation to the patient's anatomy automatically detected by the system and visualized in the User Interface?
5.2.9 | Does this apply to all surface coils?
5.2.10 | Is it visualized at the scanner what coils are plugged in?
5.2.11 | What kind of coil plugging mechanism is available?
5.2.12 | Are coils available without cables?
5.2.13 | Please list these coils:
5.2.14 | If cables are used, is the plugging mechanism operable with one hand?
5.2.15 | If cables are used, are the coil cables connected to the patient table instead of the magnet front, to ensure easier handling with shorter cable lengths?
5.2.16 | Kinematic coil and software(functional joint study)
5.2.17 | Can the patient table be moved freely (horizontally and vertically) with the coil cables plugged in?
6. Patient Handling
6.1 | Patient Table Fixed or Standard
6.1.1 | Specify Vertical table movement and Max. vertical speed
a | Minimum table hight from floor | cm
b | Maximum table hight | cm
6.1.2 | Horizontal table movement
a | Maximum horizontal range of table movement: | cm
b | Maximum speed: | cm/s
c | Accuracy of repositioning: | mm
6.1.3 | Maximum patient weight not less than 220kgs (including vertical and horizontal table movements): | kg / lbs
6.1.4 | Table Operation/ Landmark patient
a | Digital display of table position:
b | Isocenter positioning just by push-button:
c | Home positioning (out) just by push-button:
d | Home positioning (up) just by push-button:
e | Automatic positioning of patient table w/o laser:
f | Manual table move in case of emergency:
g | Detachable Option
h | Table control from console (e.g. for peripheral MRA):
6.3 | Scannable range / whole body coverage
a | Is it possible to start and stop a scan at the magnet control panel?
b | Integrated infusion stand
6.4 | In-room information for patient preparation
a | In-room display of patient’s name
b | In-room information on the exam
c | In-room display of coils being plug in
6.5 | Accessories for Patient Positioning
a | All necessary accessories for patient positioning (mattresses, cushions, vacuum cushions) shall be included.
7. Physiological Signals
7.1 | Physiologic Measurement Unit (PMU)
All the trigging techniques should be standard. The following lead and probe need to be included as standard: - One set of patient wires, - One box of adult and pediatric size electrodes and - One fiber optic peripheral gating probe
7.1.1 | ECG
a | Is the ECG signal acquired from three ECG channels, to ensure optimal quality? (Vector Cardiogram, VCG)
7.1.2 | Peripheral Pulse Triggering
7.1.3 | Respiratory Triggering
a | Is respiratory triggering possible?
b | Pneumatic cushion, to be placed on the chest, for easy handling?
If a belt needs to be wrapped around the patient, please provide details.
7.1.4 | External Triggering
a | Is an interface included, for a trigger input from external sources?
b | Optical trigger output?
7.1.5 | Handling
a | Wireless trigger signal transmission for easier placement of the trigger devices:
b | State trigger devices that operate wirelessly(e.g. ECG, peripheral pulse, respiratory)
c | Are ECG, respiration, and pulse displayed at the main console?
d | Are the trigger points visualized on the display for better control?
7.2 | Is an additional display of the physiological signals in the examination room available?
Applications
8 | Neuro Imaging Applications
8.1 | Routine T1-weighted, T2-weighted and proton density weighted imaging for head and spine imaging:
a | Fluid-attenuated inversion recovery (FLAIR, 2D and 3D) imaging for complete suppression of the CSF signal:
b | SUSCEPTIBLE WEIGHTED IMAGING (SWI)
c | 3D NERVE VIEW
d | Single-shot myelographic imaging with high resolution (true 512 matrix):
e | Multi-step imaging of the whole CNS (head + whole spine) with surface coil image quality, without repositioning of the patient or coils:
f | Multi-step imaging of the whole CNS (head + whole spine) with fully automated, software-controlled table movement:
8.1.2 | Diffusion Imaging
a | Diffusion imaging with multi shot EPI, HASTE and FSE and a high-resolution (non-single-shot) technique shall be included.
b | High-resolution diffusion imaging (non-single-shot):
c | Spine diffusion imaging for differentiation between osteoporosis, tumor infiltration and radiotherapy-induced changes:
d | Max. b-value:
e | Sequence should be able to carry out slice-by-slice tuning of resonance frequency, transmitter voltage, first order B0-shim and B1-shim. For whole-body diffusion, the SliceAdjust technology helps to avoid station boundaries and apparent broken spine artifacts as well as to preserve the SNR for whole-body diffusion.
f | Calculation of ADC (Apparent Diffusion Coefficient) maps:
8.1.3 | Diffusion Real-Time Evaluation
a | A fully automated real-time evaluation package for the evaluation of diffusion-weighted images (DWI) shall be included.
b | Can ADC maps (Apparent Diffusion Coefficient) be calculated automatically in real-time from DWI?
c | Can Trace-weighted images be calculated automatically in real-time from 3 separate orthogonal DWI scans with minimal TE?
8.2 | Neuro Workflow Features
a | Is the head of the patient automatically placed at the isocenter without any laser marking required?
b | Does the software guide the user step-by-step through the standard brain examination by showing example images and given explanation in text form?
c | How many mouse clicks does the user need to adapt an examination to the patient’s condition, e.g. switch to a protocol with motion correction if the patient is moving?
d | Are the protocols automatically positioned and aligned according to the anatomically derived sagittal, coronal, and axial slices of the localizer?
e | Are clinical decision steps possible so that the user can add or eliminate protocols or groups of protocols with one mouse click?
f | How many mouse clicks does the user need to rerun a sequence with identical parameters?
9. | Angiography
9.1 | MR Angiography Applications
9.1.1 | Time-of-Flight MRA (TOF)
a | MR-Angiography using the Time-of-Flight method shall be included.
b | 2D TOF:
c | 3D TOF:
d | Sequential 2D TOF with "travelling saturation bands":
e | 3D multi-slab (several overlapping 3D slabs in a single scan for larger anatomical coverage):
f | Automatic elimination of overlapping slices for 3D multi-slab:
g | Ramped RF-pulses for 3D scans, in order to decrease saturation effects in distal vessels:
9.1.2 | Phase Contrast MRA (PC)
a | MR Angiography using the Phase Contrast method shall be included.
9.1.3 | Contrast-enhanced MRA (ceMRA)
a | Ultrafast sequences for contrast-enhanced MR Angiography shall be included. Vascular Application, (all regular Vascular imaging sequences and the advanced application need to be included as standard), all available technique should be offered as a standard or optional.
9.1.4 | Bolus Timing
a | A software package for online-visualization of a fast 2D bolus scout and a fast switch to the 3D ceMRA scan shall be included.
10. | Cardiac Morphology
a | Cardiac Morphology Applications
b | Is a step-by-step procedure provided for fast cardiac view creation (all relevant scouts)?
c | Graphical ‘3-point planning' of cardiac views:
d | Dark-blood preparation pulses, for optimized myocard-blood contrast:
e | Breath-hold techniques with high contrast:
10.1.3 | Cardiac Function
a | The assessment of ventricular function and wall motion shall be included.
b | Spoiled GRE technique (FLASH, Spoiled GRASS, T1-FFE) for cardiac function:
c | Gradient Echo with full transverse rephasing (TrueFISP, Balanced FFE, FIESTA) for cardiac function:
d | Automatic adjustment of the acquisition window to the current heart rate:
e | Graphical display of data acquisition window (e.g. TR) with respect to the ECG signal:
f | Can all these techniques be performed during breath-hold in order to examine the whole heart in a short time without artifacts?
g | Parallel Imaging (T-PAT, iPAT, SENSE, ASSET) for highest temporal and spatial resolution:
h | Radial scanning for real-time cine studies:
i | Real-time cine TrueFISP/FIESTA/B-FFE for non-triggered, free breathing acquisitions:
j | Is echo sharing and retrogating supported?
k | Are stress/rest examinations supported?
l | MyocardiaL Tagging
m | Look Locker for optimum Inversion Delay
n | Provides single breathhold, multi-echo, ECG-triggered acquisitions to provide T2*, R2*, T2 and R2 maps for assessment of myocardial tissue characteristics
o | Are stress/rest examinations supported?
p | Can arbitrarily oriented slices (short axis AND long axis simultaneously) per heart beat be acquired?
q | TI scout for optimization of contrast between infarction and normal myocardium:
q | 2D and 3D 'delayed enhancement' techniques with segmented IR TrueFISP/FLASH sequences, for the assessment of myocardial viability:
s | 2D 'delayed enhancement' techniques with segmented IR TrueFISP/FLASH sequences, for the assessment of myocardial viability:
t | First pass myocardial perfusion imaging with multiple slices per heart beat on the basis of fast GRE or SSFP techniques?
u | Phase-sensitive IR TrueFISP/FLASH for “Delayed-Enhancement“ studies with optimized contrast; no need to adjust the inversion time TI?
10.1 | Flow Quantification
a | A comprehensive package for flow quantification shall be offered as an option.
b | ECG-triggered phase contrast technique:
c | Segmented 2D GRE with echo-sharing, to increase the temporal resolution:
d | Breath-hold techniques for flow quantification:
e | Retrospective reconstruction algorithms for full R-R coverage in combination with arrhythmia rejection:
10.2 | Liver Imaging
a | Liver imaging application: triggering techniques and new strategies for motion control and fast imaging including MRCP
b | Quantitative MR for hepatic diseases assessment/Iron Quantification/T2 star Mapping
c | Mention all available sequences.
d | MR elastography
10.3 | Oncology Application
a | In addition to all regular imaging sequences, the following advanced application need to be included as standard:
b | Whole Body screening with T1 sagittal, T1 Coronal, T2 Coronal, STIR and Diffusion Weighted Imaging, DWIBS
c | Whole body imaging using integrated body coil along with automated table stepping
11 | Breast Imaging
a | Suitable techniques to perform MR Breast Imaging and evaluation shall be included, in combination with a dedicated breast array coil.
b | Availability of fast 3D Gradient Echo sequences for dynamic contrast studies (e.g. 64 slices with 256² matrix in < 1 minute):
c | T2-weighted imaging with fat saturation (examination of cysts):
d | 'Silicon-only' sequences for the examination of implants:
e | Fully automated calculation of parametric maps for dynamic breast examinations (e.g. TTP, PEI):
f | Breast coil compatible with Parallel Imaging:
12 | Orthopedic Imaging(MSK)
12.1 | Orthopedic imaging shall be included with joints cartilage assessment/Mapping
12.2 | Metal artifact reduction
12.3 | mDixon Software
13 | Whole-Body Imaging
13.1 | Patient table for Whole-Body Imaging
13.2 | Does the table shift range allow for whole-body examinations without patient re-positioning?
13.3 | Can whole-body examinations be performed with high-SNR surface coil quality, without patient re-positioning?
13.4 | Max. scan range (without re-positioning of the patient, with high-SNR surface coil quality for the whole body):
13.5 | Whole body diffusion with automatic stiching
13.6 | Describe the technical solution for whole-body imaging:
13.7 | Whole-Body Applications
13.8 | Whole-body MR angiography (with high-SNR surface coil quality, without patient re-positioning):
13.9 | Whole-body metastases evaluation (with high-SNR surface coil quality, without patient re-positioning):
13.1 | Can Parallel Imaging be applied throughout the whole body?
13.11 | If not, state restrictions, i.e. for which body parts in a whole-body examination Parallel Imaging is not possible:
14 | Parallel Imaging
14.1 | Routine Parallel Imaging-A comprehensive package for Parallel Imaging shall be included.
14.2 | An advanced package for Parallel Imaging shall be offered, with the functionality to enable '2-dim' Parallel Imaging, i.e. simultaneous PI in phase-encoding and 3D-encoding directions with 3D scans.
14.3 | Simultaneous Parallel Imaging in 2 dimensions possible:
14.4 | Is a high-speed Parallel Imaging technique available?
14.5 | Is this Parallel Imaging technique compatible with the standard coil set?
14.6 | Explicitly state possible restrictions, i.e. coils that are not compatible with Parallel Imaging:
14.7 | Image-based algorithm for Parallel Imaging (e.g. SENSE):
14.8 | k-space-based algorithm for Parallel Imaging (e.g. SMASH):
14.9 | Is a separate measurement for coil sensitivity calibration necessary?
14.1 | What is the time needed for the coil sensitivity calibration?
14.11 | Is a new separate calibration measurement necessary if the patient and/or the patient table moves?
14.12 | Does the software support the user in finding optimal parameter settings for Parallel Imaging?
14.13 | Which parameters are taken into account by this online help function?
14.14 | Advanced '2-dim' Parallel Imaging
14.15 | Spectroscopy
14.16 | functional MRI (fMRI)
15 | STANDARD MULTI REGION Pulse Sequences . Each vendor to give his specific sequences.
15.1 | Spin Echo
15.2 | Fast Spin Echo
15.3 | Gradient Echo
15.4 | Inversion Recovery
15.5 | Phase Imaging
15.6 | Multislice/Multiecho
15.7 | 3D Volume Imaging
15.8 | Rectangular FOV
15.9 | Presaturation with auto-shim capability for fat suppression
15.1 | Minimum of 256 echo train length
15.11 | Fast Inversion Recovery
15.12 | Fluid Attenuated Inversion Recovery (FLAIR)
15.13 | Fast FLAIR
15.14 | Short Tau Inversion Recovery (STIR)
15.15 | Fast STIR
15.16 | Fast Gradient Echo
15.17 | MR Fluoroscopy
15.18 | 2DFT
15.19 | 3DFT
15.2 | FatSat
15.21 | Magnetization Transfer Contrast (MTC) and Ramped-RF (TONE)
15.22 | Motion Artifact Suppression Technique (MAST) /Motion correction
15.23 | No Wrap
15.24 | Phase Wrap
15.25 | Flow Compensation
15.26 | Gating to include peripheral pulse gating
15.27 | Echo Planar Imaging(EPI)
15.28 | Breathhold Cine
15.29 | ECG Triggering
15.3 | ECG gating
15.31 | Retrospective Gating
15.32 | T1 – FLAIR
15.33 | T1-weighted Phase Sensitive Inversion Recovery
15.34 | Double inversion recovery imaging
15.35 | balanced steady state free precession Spoiled Gradient Recalled
15.36 | Dual Echo Fast Gradient Echo
15.37 | IVIM and Kurtosis DWI
15.38 | Arterial Spin labeling technique
15.39 | Amide Proton transfer imaging (APT)
15.4 | 0 TE Options
15.41 | Bold imaging (fMRI)
15.42 | CSF study
15.43 | Dixon
16 | Image Processor
16.1 | The image processor shall have hard disks dedicated to raw data ≥600GB; main memory ≥ 128 GB.with fast high end multitable processor
16.2 | Specify Image reconstruction speed in FFT/s or images /s.
17 | Computer System
17.1 | The computer system shall have a minimum of 32GB main memory (RAM)
17.2 | The computer system shall be of an open architecture design to allow for upgradability.
17.3 | System shall be able to communicate, send and receive images to a PACS, HIS, RIS and/or teleradiology system or Any other applicable devices
17.4 | Enable remote service support
18 | Image Storage
18.1 | Multiple Hard Disk drives of large capacity ≥ 2 TB for System Software and data storage. It shall have image storage capacity of at least 400,000 images or higher at 256x256 image matrix.
18.2 | CD-DVD RW for archiving and storage of images
18.3 | Compliant with all stand DICOM Specifications
18.4 | System should be cybersecure and patient data protected with the latest preventive patches.
19 | Accessories and other requirements
19.1 | MR compatible fire extinguisher -2 NOS
19.2 | Automatic Door Mounted metal detector (Provide Qoute also)
19.3 | Handheld metal detector -2 NOS
19.4 | Non-Magnetic Wheel Chair--1 NOS and Non-Magnetic Stretchers --1 NOS
19.5 | Phantoms: For Image Quality audits and monitoring./ACR PHANTOM
19.6 | Hearing Protection & Audio: MRI earplugs, specialized headphones, and two-way intercom systems
19.7 | MR Compatible patient Music system +CINEMA VISION SYSTEM
19.8 | CCTV SYSTEM WITH TWO WAY PATIENT with Camera Color and Zoom ( SCANNER ROOM) - OPERATOR ( CONTROL ROOM)
19.9 | IV Poles & Infusion Stands: Non-ferromagnetic height-adjustable stands
19.1 | MR compatible Cabinet in the room for coils and accessories, Positioning Aids: Acoustic and non-ferrous foam pads, wedges, and straps etc 2 Set specialized patient support accessories.
19.11 | MR Compatible Dual-Head Contrast InjectorAutomatic pressure Injector
20 | Patient Monitoring requirements
20.1 | Mini Anesthesia Equipment MRI-compatible to handle sedated/GA patients, with non-magnetic ventilators Suitable for 3T scanner.
20.2 | RI-compatible patient monitor for 3T MRI environments.
20.3 | Color TFT display, minimum 12-15 inches.
20.4 | Monitoring parameters: ECG, Heart Rate, SpO₂, NIBP, Respiration, EtCO₂, and Temperature (optional).
20.5 | MRI-compatible ECG and SpO₂ sensors with artifact reduction technology.
20.6 | Automatic, manual, and interval NIBP measurements.
20.7 | Audible and visual alarms with adjustable limits.
20.8 | Rechargeable battery with a minimum 4-hour operating time.
20.9 | Data storage and trend review capability.
20.1 | Supplied with MRI-compatible accessories (ECG cable, SpO₂ sensor, NIBP cuff, EtCO₂ module, charger, and transport cart).
21 | AI Reconstruction
21.1 | AI Capabilities | Vendor-native deep learning-based reconstruction integrated into clinical workflow. | Acceleration, noise reduction, improved sharpness and spatial resolution.
21.2 | Clinical Coverage |  | Applicable to Neuro, Spine, MSK, Cardiac and Body imaging.
21.3 | Sequence Coverage |  | Supports T1, T2, PD, DWI, motion correction, implants, free-breathing.
21.4 | Evaluation Criteria |  | Additional AI advantages will be scored.
22 | Training
22.1 | Onsite basic application training for all radiographers as needed
22.2 | One week Onsite intermediate application training for all radiographers
22.3 | One week Onsite Advanced Cardiac and other special application training for senior radiographers
22.4 | One week application training upon request
22.5 | Service and maintenance training for biomedical engineers.
23 | Trade-in (If relevant)
23.1 | Supplier should offer a Trade-in price for the following existing system:
23.2 | DESCRIPTION MANUFACTURER MODEL YEAR OF PRUCHASE
23.3 | MAGNETIC RESONANCE IMAGING SIEMENS, GERMANY MAGNETOM AVANTO-FIT UPGRADE 21/07/2011(Upgraded in 2018)
24 | Quality and Patient safety
24.1 | Unit must be FDA / CE certified
24.2 | Unit must comply ISO 13485 Quality management system for medical device or equivalent .
24.3 | Unit must comply with Electrical safety test IEC 60601-1,-2,-18 or equivalent
24.4 | Quality control ,Quality assurance and acceptance test should be done by the company, QC tools for performance verification done by End-user, if required.
26 | Consumables
26.1 | Should be Supplied with all standard consumables.
27 | Accessories
27.2 | Should have All standard Accessories must be included as per manufacturer .
28 | Available Additional Option
No. | OPTIONS DESCRIPTION |  |  | REQUIRED OR NOT REQUIRED
29 | Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Complete on-site (off-site, if necessary) training for Biomedical Engineering Staff (Service Training) and End user training for operators and Physicans is required.
3 | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS ,FDA Approval ,CE Mark
4 | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
5 | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
6 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
7 | Continuous Update and Upgrading of the system and all provided software with the system including and limited to the post processing and DICOM software, during the warranty (any hardware upgrade needed for such software upgrade should be included and provided with no additional cost by the vendor). It is expected from the vendor to always install the latest software versions within 4 months of its official release.
8 | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
9 | All coils and MRI accessories, helium and Cold head, shillers should be full covered and included during warranty and maintenance contract periods.
10 | Maximum time needed for parts ordering and replacement (is calculated as a system down time )
11 | Maximum response time for system faults or breakdown calls (max. time for the engineer to be at the site) to be provided.
12 | Warranty must be minimum 3 years comprehensive from date of installation

## MRI Civil & MEP Works
Civil and MEP works for New MRI 3T
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Civil and MEP works for MRI 3T | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No
All Pre-installation requirements should be submitted, including suggested room Design and dimensions
1 | All room preparation and system installation requirements shall be clarified and agreed upon after the vendor's initial site visit. A complete room layout shall be provided.
2 | The scope shall include the proposed design drawings and the as-built drawings to be submitted upon completion of the project.
3 | Modern and nice interior design with decorative works shall be proposed through 3D Design and included as part of the project scope of work.
4 | Demolition and Mobilization works should be included
5 | Mobilization works, including wall opening and reinstatement/closing works required for the delivery and installation of the 3T MRI machine, shall be included in the scope of work.
6 | The RF cage shall be designed, supplied, and installed in accordance with SHA standards and requirements, with consideration for long-term operational reliability, durability, and maintenance. The proposed RF cage shall include all required components, such as RF doors, RF viewing window, and related accessories, along with complete technical specifications and details.
7 | The maximum expected acoustic noise level outside the MRI room shall be considered. The proposed MRI room door shall have proper soundproofing performance, and its technical specifications for acoustic noise reduction shall be provided.
8 | RF cage Testing and Commssioning Certification should be provided
9 | Partition and ceiling works shall comply with UHS/SHA standards and be fully compatible with the 3T MRI system. The proposed technical specifications and details shall be submitted for review and approval.
10 | Medical-grade flooring shall be provided, Simmilar to exsisting The proposed flooring specification, material details, and floor layout design shall be submitted for review and approval.
11 | Removing of existing flooring & levelling with self-levelling screed for complete area.
12 | Proposal for the chiller area foundation, including required barriers, shall be provided. A site survey shall be conducted if the existing MRI chiller area is not sufficient.
13 | Power Supply Capacity and Specifications
14 | Design of Power Layout
15 | Proposal for Supply and Instation of Electrical works. Socket and switches should be via brand and of same type used in UHS premises.(All wires will be concealed and in conduct pipes and the wiring for electrical and network cable will be commensurate with the existing and MRI Compatable). Confirm the load requirement with existing Available Load, if not it should be included in the proposal.
16 | The lighting should be all LED and should support energy conservation and MRI 3T Compatable
17 | HVAC- Study the current layout and Propose the modifications if required to meet the SHA Standards related to the RH and Temprature in the proposed area.
18 | Fire Alarm and Fire Fighting modification works to be proposed as per Civil Defense Standards and SHA if requireed any changes in proposed area.(FM200 - Sprinklers) etc….
19 | Works related Shifiting of current Bed head units to proposed location including Medical Gas Points and Nurse call system shall comply with SHA standards.
20 | New doors, including all required door accessories such as door stoppers, door protection, handles, and locks, shall be provided. Door locks shall be compatible with the hospital master key system.
21 | CBS System should be provided in every Area, IPC Connection FP200 4mm should be used.
22 | Sliding Door ( Dorma or , GEZE) Brand
23 | Access Control system should be integrated with the existing Access control System.
24 | Decorative Glass Picture panel in roof and Walls should be provided in MRI Room.
25 | Wall protection and corner protection should be considered.
26 | Vibration and Acoustic requirement should be fulfilled in accordance with SHA requirement .
27 | Any additional requirements related to the MRI pre-installation works, as per manufacturer recommendations, applicable standards, and hospital requirements, shall be considered and clearly included in the scope of work.
28 | Weight load assessment, slab reinforcement test, and required testing for the MRI installation location shall be provided. Any required strengthening works shall be considered and included if necessary.
29 | All other requirement that will be required for the completion of replacement of MRI scope. Should be considered .
Available Additional Option
No. | OPTIONS DESCRIPTION |  |  |  | REQUIRED OR NOT REQUIRED
1
Important Notes & Obligatory Conditions
1 | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void
2 | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Equipment. Two Copies each and one Soft Copy for each.
3 | It is mandatory for the Contractor to submit End-of-Sale, End-Of-Support, and End-Of-Life for each individual hardware component - Proof documents from the manufacturer to be attached with the proposal.
4 | Proposal Should be submitted after having site survey and comparison with the current As built
5 | All The works Should Comply with Standards of Civil defence, SEWA and MOH
6 | Contractor should discuss the final technical proposal with the technical team before submission. Proposal should include ongoing warranty, support.