## Stainer
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Fully Automated Dual Stainer for Gram stain and Acid fast Stain | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the fully automated Dual Stainer for Gram stain and Acid fast Stain and optional items should be mention clearly.Please specify the details clearly and all the requested specs information must be well supported with reference documents.
2 | Fully automated workflow system should include staining
3 | Flexible staining protocols; pre-set customized programmes.
4 | minimum 20-30 slides per cycle, maximum time requirement for one cycle 10 min
5 | Minimal reagents requirement ,with no or simple manual reagent preparation steps
6 | Easy maintenance and trouble shooting
7 | user friendly menu options with easy swtiching between the stain programs
8 | Slide spraying method not dipping or any other advanced technique
9 | Minimum on board Reagent stability should not be <30 days
10 | Reagent supply should be smooth. Lot to Lot variations to be minimal
11 | LCD/LED screen display
12 | Automatic cleaning after each cycle
13 | Slide heating and drying
14 | built-in heating plate for AFB stain (ziehl-neelson stain)
15 | Availablity for demo
16 | Quality requirements
16.1 | FDA/CE mark approval
16.2 | Atleast 2 Quality controls for each stain (Gram stain: Gram positive/Gram negative) (Acid Fast stain: positive AFB/Negative AFB)
16.3 | Number of customers used in CAP accredited labs in UAE / ME Region
16.4 | CAP or other external quality control program enrollment
17 | Sample Consumption
Number of Gram stain slides stained - 450-600/Month , AFB- 15-25/month
18 | Accessories
Supplied with all standard accessories.
19 | Consumables
Supplied with all standard consumables.
20 | Available Additional Option
No. | OPTIONS DESCRIPTION
21 | Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and End user training for lab staff and later as needed .
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS,CAP
D | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
E | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
F | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
G | Warrenty : Comprehensive minimum 5 Years warrenty .

## Autoclave
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| Autoclave: Fully automated vertical sterilization system | 1 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the Autoclave: Fully automated Sterilization system and optional items should be mention clearly.Please specify the details clearly and all the requested specs information must be well supported with reference documents.
2 | Fully automated strerilization system
2.1 | Flexible sterilization (liquid, solid, semi-solid); pre-set customized programmes.
2.2 | Autocalve temperature range from 100 -135 °C, with maximum 0.26 mpa pressure range
2.3 | After sterilization Agar temp must be maintained to prevent clog for atleast 10-12hrs
2.4 | Easy maintenance and trouble shooting
2.5 | Free standing, vertical design
2.6 | LCD/LED screen display, user friendly menu options
2.7 | Minimum size 85 liters, accomodating (500-700 perti plates)
2.8 | Automatic cleaning after each cycle
3 | Safety features
3.1 | Temperature protection
3.2 | Door lock
3.3 | Pressure safety
4 | Unit must be availablity for demo
5 | Quality requirements
5.1 | FDA/CE mark approval
5.2 | Compatible with Biological indicator and chemical indicator as QC material
6 | Sample Consumption
500-700 PLATES
7 | Accessories
Supplied with all standard accessories.
8 | Consumables
Supplied with all standard consumables.
9 | Available Additional Option
No. | OPTIONS DESCRIPTION
Important Notes & Obligatory Conditions
A. | Completeing this compliance sheet is manditory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) training for Biomedical Engineeing Staff (Service Training) and End user training for lab staff and later as needed .
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS,CAP
D | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
E | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
F | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
G | Warrenty : Comprehensive minimum 5 Years warrenty .