## ECG machine tender spces
Indicative Specifications / Evaluation Form
Item No. | Item Description | Qty. | To Be Filled By Vendor |  | To be filled
| 12-Channel- Electrocardiograph (ECG) with trolley | 20 | Comply | Supporting Documents Ref. | by Hospital
| Specifications |  | Yes/No (If Yes Specify)
1 | Technical offer must be clear in terms of standards , which will be offered with the ECG Machines and Optional features and items . Please specify the details clearly and requested specs/info must be well supported with reference documents.
2 | The offered equipment shall be a 12-lead diagnostic electrocardiograph suitable for adult and pediatric ECG acquisition. The ECG shall provide simultaneous acquisition and display of all 12 standard leads.
3 | Please specify make, model and origin of the machine.
4 | 12-lead ECG acquisition with simultaneous sampling.
5 | The ECG shall be equipped with a 12-lead ECG acquisition module capable of simultaneous acquisition of all standard 12 ECG leads. 15/16-Leads ECG or higher is preferable (Optional). Each individual ECG lead (limb leads RA, LA, RL, LL and individual chest/precordial leads (V1–V6). wire shall be independently detachable and replaceable in case of damage, breakage, connector failure or other defect.
6 | Minimum sampling rate: 500 samples/second/channel or better.
7 | 12-bit or higher A/D resolution.
8 | Automatic baseline stabilization and artifact/noise reduction.
9 | AC/interference filtering: 50/60 Hz.
10 | Baseline wander and muscle-artifact filtering.
11 | Input protection against defibrillator discharge.
12 | Pacemaker pulse detection capability.
13 | Lead-of/Lose contact detection.
14 | ECG acquisition time -minimum 10 seconds of diagnostic-quality data per recording.
15 | Automatic gain selection and manual gain settings.
16 | Sensitivity: preferably 5, 10 and 20 mm/mV with automatic setting.
17 | Paper speed: preferably 25 and 50 mm/sec, with additional selectable speeds where available.
18 | High-resolution colour LCD/TFT display. Display size minimum 12 inches or better bigger size.
19 | Display of patient information, heart rate, lead status, gain, filter settings and recording status.
20 | Automatic heart-rate calculation.
21 | Automatic measurement of PR, QRS, QT/QTc, Axis and other relevant ECG intervals.
22 | Automatic/Manual ECG interpretation/analysis. (ADULT & PEADIATRIC Criteria)
23 | Detection of common arrhythmias and ECG abnormalities.
24 | Pediatric ECG interpretation capability is preferred.
25 | QT/QTc calculation using a recognized algorithm.
26 | Patient demographic data entry through alphanumeric keyboard/touchscreen. & Worklist & Barcode / QR Code Scanner | 20
27 | Minimum internal storage: 800-1200 ECG records or higher. Internal memory shall retain ECG data in the event of temporary battery/power failure.
28 | Minimum battery operating 8 hours of normal clinical use or minimum 200 ECG recordings on a full charge, whichever is applicable to the manufacturer's tested specification.
29 | Integrated thermal printer A4 Paper size.
30 | Minimum printing capability: 12-lead ECG report.
31 | Standard paper size appropriate for the proposed printer.
32 | Printing speeds of 25 and 50 mm/sec.
33 | Automatic printout of patient demographics, ECG measurements and interpretation.
34 | Printing of 3×4, 3×4+1 rhythm lead, 6×2 or equivalent configurations is preferred.
35 | ECG records shall be retrievable by patient ID, name, date or examination number.
36 | Stored ECGs shall be exportable in a standard digital format.ECG must support Latest SMB(3.1.1) Protocol.
37 | USB connectivity for data export and software/service functions
38 | The ECG machines MUST have integrated/built-in Wi-Fi.
39 | Allow wireless transmission of ECG reports and ECG data to a hospital information system HL7/server.The possibility of integrating with the currently available cardiology PACS software (Philips CVIS).
40 | Any required HL7/DICOM interface license, software, middleware and activation required for operation shall be included as standard.
41 | All 20 ECG machines shall be connected to a centralized ECG management server/Central station /middleware through the hospital network.
42 | Authorized users shall be able to review ECGs from a central workstation/station without requiring physical access to the individual ECG machine
43 | ECGs shall be searchable using: Patient ID/MR,Patient ID/MRN,Patient name, Accession/order number,ECG machine/device, Department/location and Physician/user, where applicable.
44 | The system shall support network-based access from any authorized department/workstation within the hospital.
45 | Any license, middleware, interface engine, connector or software required to achieve the specified HL7 integration shall be included in the supplier scope of work .
46 | Lightweight and compact design suitable for clinical rooms.
47 | Easy to clean surfaces and disinfectable parts.
48 | Unit must be antimicrobial protection
49 | Power supply: 220–240V, 50/60 Hz
50 | Please specify the weight and unit must be provided with good quality Trolly with castors .
51 | The Machine shall support seamless integration with the Hospital Information System (HIS) to enable automated synchronization. Below are the technical Integration Requirements
52 | The proposed solution shall support standard HL7 Version 2.5 messaging for integration with the Hospital Information System (HIS).
53 | The supplier shall be fully responsible for designing, implementing, testing, and commissioning the HL7/DICOM integration with the Hospital Information System.
54 | All software, interface components, middleware, interface engines, licenses, and configuration required for successful integration shall be included within the supplier's proposal.
55 | The supplier shall provide on-site technical support during installation, testing, Go-Live, and stabilization until successful completion of the integration project.
56 | The supplier shall provide complete technical documentation.
57 | The supplier shall ensure that the integration complies with international HL7 standards and healthcare interoperability best practices and shall work collaboratively with the Hospital Information System vendor to achieve successful end-to-end integration.
58 | The supplier shall provide all software licenses, interface licenses, middleware licenses, communication licenses, APIs, drivers, connectors, and any other software components required for successful integration with the Hospital Information System (HIS).
59 | All licenses required for the development, testing, production, and future operation of the integration shall be included in the supplier's proposal without any additional cost to the hospital unless explicitly stated in the commercial offer.
60 | ECG Cable with IEC (International Electrotechnical Commission) Color coding Standard, with Alligator/Grabber type (Not Snap Type).
61 | Trade-in offer -Supplier should offer a Trade-in price for the following existing system(17 QTY ) :
1 | DESCRIPTION MANUFACTURER MODEL Qty
2 | ELECTRO CARDIOGRAPH- 12 CHANNEL GE HEALTHCARE, USA MAC 5500 12
3 | ELECTRO CARDIOGRAPH- 12 CHANNEL PHILIPS, GERMANY/USA TC-20/TC-35/TC70 |  | 5
62 | Quality and Patient safety
63 | Device should be CE and FDA approved
64 | Offered System must comply with Electrical safety IEC 6060-1
65 | Device must be manufactured on quality standards .Medical device Quality management system , ISO-13485 or equivalent.
66 | Electrical leakage and overload protection
67 | Accessories
67.1 | All necessary accessories
68 | Consumables
68.1 | Supplied with all standard consumables-Disposable
69 | Available Additional Option
No. | OPTIONS DESCRIPTION
70 | Important Notes & Obligatory Conditions
A. | Completing this compliance sheet is mandatory. All the information to be provided by the vendor in the compliance sheet must be 100% true. Any false information given will result in an immediate disqualification and the offer shall be considered null and void.
B. | Complete on-site (off-site, if necessary) end user training for Pharmacy doctors and other associated staff and Service training for Biomedical Engineering Staff is required.
C. | Equipment Meets OR Exceeds IEC - 601, ISO, ANS, BS, JIS
D. | Rated Voltage: 240V(+/- 5% ) 228-252 Volts single Phase OR 415 ( +/- 5% ) 395 - 435 Volts three Phase, 50 Hz .Three core cable and the casing must be earthed. Three pin electrical plug.
E. | A Price List For Spare Parts: Must be provided to cover at least 7 years of operation and maintenance of the equipment supplied.
F. | Operation And Service Manuals: All Manuals must be supplied as " Standard Requirements of Technical Documents " for all Medical Equipment. Two Copies each and one Soft Copy for each.
G. | Installation: All equipment must be installed by the manufacturer's engineer(s) or those approved from the Hospital Biomedical Engineering Department.
H. | Warranty : Comprehensive minimum 3 Years warranty .