## Table 2
RFP Reference: UHS/ENG/TENDER/040/2026 29th September 2026 RFP Closing Date: 13th October 2026

## Table 3
No. | Description
1 | ICU Expansion Project

## Table 4
University Hospital Sharjah. (UHS) Management has decided to invite vendors for a Request for Proposal (RFP). You, as a vendor are requested to participate in the RFP process by submitting your offer to provide the services as described in this document.

## Table 5
The RFP should comply with the following terms & conditions:

## Table 6
1. The proposal shall be clear, informative & include as per the requirement described in the RFP. 2. The financial offer should be on your company letterhead containing the authorized signatory and must be sent to the attention of the Director of Finance, University Hospital Sharjah, PO Box 72772, Sharjah in a sealed document.

## Table 7
3. The price quoted is as mentioned in the technical requirement listed below (RFP) to UHS.

## Table 8
4. All deliveries should be made for the ordered quantity in full, without partial shipments, to our Main Warehouse, located in UHS vicinity or as specified on the Purchase Order/ Contract. Failure to comply with the agreed delivery schedule or any shortfall in quantity may result in penalties or contract termination, as per the Purchase Agreement Terms and Conditions.

## Table 9
5. As a part of the RFP document, the Vendors are requested to provide their valid Trade License, Name, and Designation of the Managing Director/General Manager/Sr. Manager having the authority to bind their company for the business relationship. Also, the vendors are required the provide licenses, certificate confirming that the vendor is legalized to operate the proposed business activity. As well as the following documents:

## Table 10
a) Valid Trade License b) Updated Company Profile c) Tax Registration Certificate (TRN) d) Full Company Address & Contact Details e) Memorandum of Association (MOA) and Power of Attorney (POA) for authorized signatory (if applicable) f) An official Authorization Letter/Agency Certificate, confirming the vendor's legal authorization to supply the specified items on behalf of the manufacturer or principal company. g) Any additional approvals or compliance documents mandated by government authorities for the supply of the specified equipment. h) Non-Liability Letter and Legal Clearance Confirmation. i) Insurance Policies (General Liability, Professional Indemnity, etc.).

## Table 12
j) Declaration of No Ongoing Legal Disputes. k) Vendor Code of Conduct Acknowledgment. 6. Standard payment terms are 90 days from the date of completion of delivery of supplies/ services or as specifically agreed in purchase contract/ agreement.

## Table 13
7. Any delays or non-conformance may result in the termination of Services agreement and/or imposition of penalty for delayed services as per the Services Agreement terms and conditions. A performance bond may be required to ensure commitment to the agreed timelines and quality standards.

## Table 14
8. The proposed services shall be evaluated & approved by UHS’s before confirmation. Once the agreement is signed off, the services will have to correspond to the required services with specific time-frame, and as originally proposed, agreed and any deviations shall be considered a breach of service contract/agreement.

## Table 15
9. The specified brand and manufacturer must remain unchanged throughout the contract period unless otherwise approved by UHS in writing.

## Table 16
10. UHS will be constantly evaluating the compliance of Contract/Agreement Terms and consistency in performance of the services throughout the duration of the agreement.

## Table 17
11. Vendors are required to submit regular progress reports at agreed intervals detailing progress, challenges, and actions to address any delays or issues. Should Vendors not meet the requirements of UHS, UHS reserves the right to terminate the agreement if the vendor is not able to rectify during the time allotted by UHS’s representative.

## Table 18
12. Vendor Contact details (landline, mobile, emails) of the authorized representatives should be mentioned.

## Table 19
13. Tenders should be submitted in two sealed envelopes and submitted to the Administration Office Finance Department- UHS:

## Table 20
a. The Technical Specification details (PLEASE DO NOT INDICATE ANY FINANCIAL VALUE IN THIS). If requested for additional clarifications and details these need to be submitted to (Administration Office Finance Department- UHS).

## Table 21
i. The offer should must conform to the RFP Document as per the attachment. ii. The offer shall be submitted (hard copy and soft copy saved in USB). iii. Reference Project/Hospital where similar work was performed.

## Table 22
b. The Financial Offer addressed to UHS’s Director of Finance, with tender reference.

## Table 23
All above documents should be submitted before the tender expiry date, all documents submitted after the expiry date will not be accepted.

## Table 24
14. UHS shall have no obligation to accept any tender proposal submitted by any vendor. UHS may at its sole discretion and without providing any reason, accept or reject any or all

## Table 26
proposals, in whole or in party. Such rejection shall not give rise to any claim, liability, or cause of action of any king by the vendor against UHS. 15. Submission of a tender proposal shall not create and agreement, legal or other relationship between the vendor and UHS. No vendor shall acquire any rights, interests, or claims against UHS by submitting a proposal, participating in the tender process or relying on any communications related to the tender. 16. In the event UHS accepts a tender proposal of a vendor, the parties agree any such tender award will be subject to a Services Agreement and separate agreement outlining the specific terms and conditions of the project and services agreed. 17. All costs, expenses or losses incurred by the vendor in connection with the preparation, submission of presentation of its proposal shall be borne sole by the vendor. UHS shall have no liability, under any circumstances to reimburse, compensate or indemnify the vendor whether in part or in whole for such costs or expenses. 18. The vendors acknowledge and agree that they have not relied on any statement, representation, warranty, or promise made by UHS, whether oral or written, in preparing their tender proposal and all decisions and judgements regarding the submission of their proposal are made at their own discretion and risk. 19. UHS may at any time without liability, amend, suspend, or withdraw the tender invitation in whole or in part. UHS may also request additional information, clarifications or documents from any vendor and may reject any proposal that is incomplete, unclear or does not comply with the tender requirements as outlined in this document. 20. Quality, Price, and sale services are combined parameters for tender evaluation. Once a vendor has been selected, a negotiation period will follow to allow both parties to review the agreement terms thoroughly. This will ensure that all deliverables, KPIs, and expectations are clearly outlined before the final agreement is signed. 21. . Vendors must submit a risk management plan, identifying potential risks to the project, such as security and confidentiality breaches, system failures, and disruptions to delivery schedules. Vendors should outline how they intend to address these risks, including their disaster recovery and business continuity plans. 22. Vendors are encouraged to adhere to ethical practices and sustainability standards in their operations. This includes providing energy-efficient equipment and adopting environmentally friendly practices in their supply chain and delivery. 23. The Vendor, its employees, its subsidiaries, and everyone who has a direct or indirect relationship with implementing and securing the works and Services included within the scope of this tender, shall be obligated to inform UHS and disclose in writing any case of conflict of interest or any private interest that has arisen, will arise, or may arise. For any transaction related to the activities of UHS, in accordance with UHS policies. 24. The vendor, its employees, and subsidiaries shall be obligated to maintain confidentiality of any data, drawings, documents, or information related to the tender - written or oral. Vendors must ensure that any data shared is protected by encryption standards and secure transfer protocols. Additionally, vendors are required to notify UHS of any data breaches immediately. Compliance with relevant data privacy regulations (e.g., GDPR, UAE Data Protection Law) is mandatory. This includes all dealings, affairs, or secrets related to UHS they may have encountered during the tender process. Vendors shall not be allowed to disclose any information related to the tender through any media outlet without obtaining prior written approval from UHS.

## Table 27
25. The copyright, rights and ownership of any documents, materials and information submitted by UHS within this tender is owned by UHS, and accordingly, these documents and materials

## Table 29
may not be copied, in whole or in part, or reproduced, distributed, made available to any third party, or used without obtaining prior written approval from UHS. If the vendor develops any custom software or systems for UHS as part of this tender, UHS will retain ownership of the intellectual property or have clear licensing terms for its continued use. All documents submitted by the UHS in connection with the request for proposals shall be returned upon request without any copies being retained by the bidder or any other person.

## Table 30
26. The vendors shall indemnify, defend and hold harmless UHS, its officers, employees and agents from and against any and all claims, liabilities, losses, damage costs, or expenses arising out of or in connection with:

## Table 31
a. the vendors participation in the tender process b. any errors, omissions, misrepresentations or inaccuracies in the proposal c. any breach of the vendors’ obligations under this tender invitation

## Table 32
27. To the maximum extent permitted by law, UHS expressly excludes any liability for: a. Any direct, indirect, incidental, consequential or special losses b. Loss of profits, revenue, goodwill or business opportunities c. Any claims by third parties arising from a vendor’s proposal d. Any loss or damage caused by errors, omissions or delays in the tender process 28. This tender invitation and all matters relating to it shall be governed by any construed in accordance with the laws of the United Arab Emirates. The competent courts of Sharjah, United Arab Emirates shall have exclusive jurisdiction over any disputes arising from or in connection with this tender. 29. This document and clauses therein constitute the entire understanding between UHS and any vendor regarding liability, proposal submission, and the tender process. No other communication, agreement or understanding, whether oral or written shall be deemed to modify, supersede, or expand these clauses.

## Table 33
University Hospital Sharjah

## Table 35
Request for Proposal (RFP) ICU Expansion Project

## Table 36
This project involves the conversion of the existing Family Medicine Department, located on the first floor of UHS, into a full ICU department and will be opened as an extension to the existing ICU-A and HDU areas; therefore, all finishes, colors, materials, themes, and overall design shall follow the existing ICU-A/HDU theme to ensure a consistent and matching final appearance, while minor differences may be accepted where required to achieve better functionality and operational efficiency.

## Table 37
The New Proposed ICU department consists of:

## Table 38
1. 5 ICU Bays

## Table 39
2. Procedure Room

## Table 40
3. Medication Room

## Table 41
4. Meeting Room

## Table 42
5. Staff Lounge

## Table 43
6. Corridor Area

## Table 44
7. Dirty Utility Room

## Table 45
8. Male / Female Locker Rooms

## Table 46
9. Male / Female Washrooms

## Table 47
10. Nurse Station

## Table 48
Civil Scope of Work

## Table 49
All participants shall submit the BOQ and prices as mentioned below. If there are any additional items that are not mentioned in our scope, please mention them at the end of your submission.

## Table 50
1. Drawings, cleaning, temporary entrance, and temporary partitions for isolation of the area. The Contractor shall include all necessary preliminaries required for the proper execution of the works.

## Table 51
2. Demolition of the existing department and disposal of debris. Any reusable items, such as doors, signage, and MEP services, shall be carefully removed and handed over to UHS.

## Table 52
3. Gypsum board partitions up to 3 meters in height.

## Table 53
4. Cement board partitions/walls for wet areas. The cement board system shall be suitable for wet- area applications.

## Table 54
5. Marble thresholds for windows and floors as required. The marble shall match the approved design and existing department finishes.

## Table 55
6. Vinyl flooring, Gerflor or equivalent. The flooring shall be suitable for healthcare/ICU applications and installed complete with all required accessories and finishing matching existing ICU theme.

## Table 57
7. Water proofing & RAK Ceramic tiles or equivalent for wet areas.

## Table 58
8. New doors complete with all accessories, including door stoppers, locks compatible with the UHS master key system, door closers, handles, and all other required.

## Table 59
9. Medical-grade sliding glass doors for all ICU bays. The doors shall be complete heavy duty and top quality with all required accessories.

## Table 60
10. Maximum possible glass windows in front of each nurse's table, allowing the nurse to clearly see inside the ICU bay while seated outside. The maximum possible height and width shall be considered, subject to site conditions and structural requirements.

## Table 61
11. Double curtains inside all ICU bays and the Procedure Room. The curtain shall include all accessories.

## Table 62
12. 60*60 Ceiling works for the corridors and ICU bays mating exsisting height

## Table 63
13. LED acrylic picture panels at all ICU bays. The picture panels shall extend from the ceiling and reflect the 3D design as shown in the approved drawings/design.

## Table 64
14. Floor-mounted and overhead cabinets, with one side for the Dirty Utility Room, two sides for the Clean Utility Room, and one side for the Procedure Room. All cabinets shall be complete with corian countertops, shelves drawers, and required accessories as per the approved design.

## Table 65
15. Complete washroom fit-out, including all sanitary fixtures, fittings, and accessories.

## Table 66
16. Staff room table as per the drawing, with Corian-finish countertop.

## Table 67
17. Strip lighting to be provided underneath each overhead cabinet.

## Table 68
18. Fansy Nurse Station / Reception counter with Corian top finish, as per the submitted 3D design.

## Table 69
19. Full corridor wall protection up to the handrail level, following the existing UHS ICU theme and design. The wall protection shall match the existing ICU-A and HDU areas as closely as possible.

## Table 70
20. Hygienic paint to all remaining areas of the department. The paint color and finish shall match the existing ICU-A and HDU theme.

## Table 71
21. Full wall guards and corner guards in all areas without existing protection, including inside the ICU bays and rooms.

## Table 72
22. Acrylic sheets behind all sinks as required.

## Table 73
23. Signage throughout the department, matching the existing UHS ICUdesign.

## Table 74
24. Sliding doors at the department entrance and the both doors for Procedure Room entrance, complete with automatic sliding door mechanisms. GEZE, Dorma, or equivalent 25. Bridge-type pendant supports shall be provided in the Procedure Room, including all required supports and structural fixing. Dynamic load testing and all necessary testing and certification shall be carried out, and all required outlets and electrical sockets shall be provided in accordance with the applicable standards and approved drawings.

## Table 76
26. The Procedure Room shall be designed to operate as a negative-pressure, positive-pressure, or neutral-pressure room, with the necessary provisions and controls to allow it to be used either as a standard room or as a Procedure Room in the future.

## Table 77
MEP Scope of Work

## Table 78
HVAC Supply and Installation 1. GI Ducting Installation:  Install GI ducting for air distribution, including both modifications and new installations, as per site requirements. 2. Volume Control Dampers (VCDs) and Fire Dampers (FDs):  Supply and install VCDs and FDs as required for the project. 3. Mineral Wool Insulation:  Install mineral wool insulation as per hospital requirements for the proposed layout to ensure proper air distribution, adhering to ASHRAE standards. 4. Ceiling Diffusers:  Supply and install powder-coated square and round ceiling diffusers. 5. Heater Batteries:  Install, test, and commission heater batteries for all areas: 6. BMS Integrated Thermostats:  Provide and commission 6 BMS integrated thermostats. 7. LTHW Piping and Valve Package:  Supply and install low-temperature hot water (LTHW) piping and valve packages for heater batteries. 8. Water and Air Balancing, Testing, and Commissioning:  Perform water and air balancing, testing, and commissioning for the renovated area. 9. Identification Stickers and Bends:  Provide and install identification stickers and bends as required. 10. As-Built Drawings:  Prepare and submit as-built drawings, incorporating existing services and structure drawings. Scope of Work: BMS Control 1. Supply and Installation of Control Equipment:  The contractor shall include the supply of all items of control equipment such as valves and thermostats, and damper motors/actuators. The BMS contractor shall also include the off-loading, positioning, and fixing of all control panels. 2. Start-Up Test and Validation:  Conduct start-up tests and validation of the BMS. 3. As-Built Composite Electric Diagrams:  Provide as-built composite electric diagrams showing interlocks between equipment furnished under this and other sections, and controls furnished herein. 4. Completion Reports:  Upon completion of work, submit reports of check-out and successful commissioning of the BMS.

## Table 79
Scope of Work: Plumbing Supply and Installation The scope of work includes, but is not limited to, the supply and installation of the following: 1. Domestic Water Supply Piping:  Install cold and hot domestic water supply piping.  Consider piping supply and return for the Ro water hymo dialysis in the bays and procedure room similar to HDU 2. Auto Sink Mixer:  Supply and install auto sink mixers. 3. Wash Basin: 4. • Supply and install high-depth, sinks inside the ICU bays, 55 cm sinks in the other rooms, and stainless-steel surgical scrub sinks for the Procedure Room, 5. Drainage Piping:

## Table 81
 Install drainage piping as required. 6. Scanning for Drainage Connection:  Perform scanning of the ground floor, lower level, and basement floor at high levels for drainage connections. 7. Core Cutting:  Conduct core cutting as per site requirements. 8. Floor Chipping for Drainage Piping:  Perform floor chipping for drainage piping for the required area 9. Pressure and Leak Testing:  Conduct pressure and leak tests in accordance with standards. 10. Identification Stickers and Bends:  Provide and install identification stickers and bends as required. 11. As-Built Drawings:  Prepare and submit as-built drawings, incorporating existing services drawings. Scope of Work: Central Medical Gas System Supply and Installation 1. Supply and installation of Central Medical Gas from 2 different AVSU panels (Medical Oxygen, Medical Air, and Vacuum 2 outlet for each) System in the proposed area as per SHA Standards. Bedhead Unit (BHU) Provide and install a complete hospital Bedhead Unit (BHU), suitable for clinical use, comprising the following:  The two outlets of each medical gas service shall be arranged such that the supply to each outlet originates from a different AVSU, providing independent supply paths.  Medical gas pipework, isolation, connections, outlet types, identification and testing shall comply with the applicable medical gas standards and local authority/hospital requirements.  6 × Twin UPS sockets (12 socket outlets total).  3 × Generator power sockets.  All electrical outlets shall be suitably labelled and segregated according to the electrical supply source.

## Table 82
2. Complete System Design and Installation:  This includes the complete detailed design, supply, erection, setting to work, testing, commissioning, and handing over of the medical gas pipeline systems to form a complete installation as scheduled and in accordance with the specifications and in full compliance with all relevant standards and the requirements of HTM -02 1. Testing Standards:  Medical gas installations shall be tested in accordance with (HTM-02) standards for medical gases installation, and Euro Pharmacopoeia for medical gas quality. 2. Medical Grade Piping:  Install medical grade piping from the existing provision near the nurse station in the pediatric ward. 3. Future Connection and Isolation Valves:  Tap-up from the existing future connection and extend for future connection with a new set of isolation valves above each room there should be a separate valve and Nest option 4. Pressure and Leak Testing:  Conduct pressure and leak tests in accordance with standards. 5. Identification Stickers and Bends:  Provide and install identification stickers and bends as required. 6. As-Built Drawings:  Prepare and submit as-built drawings, incorporating existing services drawings.

## Table 83
Scope of Electrical Work

## Table 84
Power and lighting. 1. Additional feeder (ABB) brand for the ups power , Addition of section (For Normal /emergency /UPS) in DB and cable pulling from nearby DB in Electrical room in Pediatric ward. 1. PVC and GI Conduit & cable pulling for the required power outlets and lighting as per new layout.

## Table 86
2. Lights of approved Lux shall be installed as per Clinical requirement with 50, 000 Hrs life span. 3. Cable tray and trunking to be provided as per the area requirement. 4. Earth strips be provided for all the mechanical containment and connected to main earthing system. 5. Lights in the cooridor should have dimmable system 6. 6 ups twin socket should be considered at medication room 7. Future installation of tv should be considered support and power 8. Cctv cabling only should be consider 9. Sound system and public address should be considered 10. Access control system for all rooms and entrance

## Table 87
Central Battery 1. In each separate area one light shall be connected to the Central Battery System via through SCM and SVAEL, along with programming and updating in easy check panel. 2. Civil Defense approved Exit Light shall be installed as per NFPA and Local Authority Standards. 3. Updating the Evacuation plane as per new proposed layout.

## Table 88
Fire Alarm System 1. In each separate area smoke detector shall be installed and programmed below and above the Ceiling as per NFPA and civil defense requirement. 2. Manual Call point Point and CR interfaces shall be provided as per requirement .

## Table 89
Telephone Data 1. Cat 6 UTP Schinder Shall b provided for the telephone and data network . 2. Trunking and Containment, wiring, final fix testing and commissioning from point to IDF panel and patch panel shall be provided as per requirement. 3. Labeling shall be provided to all the telephone outlets.

## Table 90
Nurse Call System

## Table 91
Containment, wiring and final fix shall be installed and programmed as per existing make ( Honeywell) including integration main system.

## Table 92
Access Control system .

## Table 93
Access control system by Honey well and shall be integrated with the existing system. Scope includes testing commissioning and programming need to be done Que System containment wiring and final fix need to be done as per stander 11. Access control system need to install as per stander and need to connect with central access control system and testing commissioning and programming need to be done and Belden cables can be used for wiring and card reader make required HID.

## Table 95
Material Specification Low Temperature Hot Water Heater Battery A low temperature hot water coil (LTHW) coil shall be provided to heat the supply air. The valve position shall be modulated when the normal operating period is signaled by the BMS and the supply fan is proven. The valve is positioned in response to a PI control signal to obtain the required air temperature.

## Table 96
Fire Dampers (F.D) 1. Whether shown on drawings or not, supply and install UL listed or BS equivalent fire dampers at all positions where ducts/builders work holes or return air transfer openings penetrate fire rated walls, floors and partitions. Sizes to match duct sizes and ratings to match the fire rating of the fire separation. Fire dampers shall conform to the requirements of NFPA 90A or BS 476 part 8 and the Department of Civil Defense. 2. Fire dampers shall be provided as noted in ducts and openings, whether shown on the drawings or not, in the following locations:  Fire Walls  Fire Partitions 3. Each fire damper shall be provided with an access door sized to permit inspection and resetting of damper and located in adjacent ductwork. Access Doors 1. Access/inspection openings shall be provided to facilitate inspection and cleaning of all ductwork systems. 2. Access doors in horizontal ductwork shall be installed at 10 meter centers, to facilitate cleaning. 3. In addition, access doors are required as follows as well as stated in other sections of the specification.  Fire/smoke dampers – one side  Heating/Cooling coils – both sides 4. Sizes of Access panels/doors shall be as HVCA TR/17, Tables 3 and 4 and shall not be less than 200mm x 150mm.

## Table 97
Flexible Ducting 1. Flexible ducting shall be provided for connections between ductwork, terminal units, diffusers and where indicated on the Drawings. It shall conform to the requirements of BS 476 Parts 6, 7 (class 1) & part 20 (15 minutes). 2. Maximum lengths shall be 1.5 meters, and all ducting shall be adequately supported to prevent oscillation and noise generation. 3. Flexible connections shall be kept as short as possible and shall be supported rigidly to prevent movement due to air flow. Ductwork Identification Markers 1. Comply with DW 144 for identification of ductwork. 2. Direction of flow markers shall be provided on all ductwork to identify the service and direction of flow. Markers shall be of triangular shape of 150 mm length of side and point in the direction of flow. 3. Direction of flow markers shall be provided that is easy 4. visible locations at approximately 3.5-meter intervals. Insulation Mineral Wool 1. Mineral pipe insulation sections shall have a nominal density of not less than 120 kg/m³. Pre-formed pipe sections shall be resin bonded with a thermal conductivity not exceeding 0.038 W/mK at 50 °C me an temperature 2. Mineral wool duct slab shall have a nominal density of not less than 45kg/m³ with a thermal conductivity not exceeding 0.04 W/mK at 50 °C mean temperature. 3. Thickness of insulation shall be in accordance with BS 5422: 2001 Piping Works Valves, cocks, air vents and pipework accessories shall be provided where indicated on the drawings and at all positions necessary for the proper working, regulation, control and maintenance of the installation with the approval. Double regulating Valve Fixed orifice double regulating valve to BS7350 Valve body material: Bronze Pressure rating: PN25 End connections: Threaded to BS 21 Specification: Fitted with two insertion points allowing quick connection; Approved Manufacturers / Vendors ;

## Table 99
Sr.No | Material Description | Make
HVAC
1 | Heater Batteries | Bissol or equivalent
2 | Duct Sensors & Thermostat | Honeywell
3 | Galvanized Iron ( GI) Duct | Nippon -Japan, Or equivalent
4 | Mineral Wool Insulation | Rock Wool -Fujairah
5 | Volume Control Dampers | Local
6 | MS Piping for Heater Battery | Suriya
7 | MI Fitting for LTHW | Crane
8 | Heater battery Valves Assembly | Crane
9 | Flexible Duct | Supaflex or equivalent
10 | Flexible Duct Rigid | Supaflex or equivalent
11 | Disc Valves | Betacad
12 | Air Diffusers | Betecad
13 | Angle/Support/Nut/Rods | N/A
14 | Control valve | ACSYS Control
15 | Thermostat Make - Alerton (Honeywell) Model - VLD-362-FF | BMTS
16 | Fresh Air Handlinh Unit Flow 1700 l/sec. with Chilled water flow on 3rd Floor Mechanical room | York
17 | Twin Exhaust fan 1700l/sec at Roof Top | Rosenberg
Sanitory/Plumbing
1 | Wash Basin | RAK-Karla
2 | Wash Basin – Mixer Auto | RAK-Kludi - Zentia
3 | Floor Drain Cover | Generic
4 | Acrylic Sheet wall protection 3 mm thick white Milky | Generic
5 | Bottle Trap with Waste | Vigga – CP
6 | Domestic Water Supply PPR Pipe | Aquatherm- Germany
7 | Soil & Waste Drainage PVC Grey | Hepworth Cory
Fire Fighting
1 | Pendant Sprinkler Head NPT Thread K-5.6, 79C Standard response Bronze finish UL /FM ( RA3415) Make Reliable Model F156 | Al Arabia
2 | Concealed Pendent Sprinkler Head NPT Thread K-5.6, 74C Quick response White cover plate finish UL /FM ( RA3415) Make Reliable Model G5-56 | Al Arabia
3 | GI Piping | Suriya Schedule 40
4 | GI Fitting | SIAM -Thailand.
Medical Gas System
1 | Bedhead Units | Dorsys-UAE
2 | Copper Piping | Lawton-UK
3 | Lockable Line Valves | Medical Grade
4 | Medical Gas Valves | Medical Grade

## Table 101
5 | Area Service Unit c/w Alarm | Beacon Medas or equivalent as per approved sample
Electrical & Low Current
1 | 2 Gang DP With Red R/R | MK
2 | 2 Gang DP With Neon | MK
3 | 1 Gang DP With Neon | MK
4 | 1 Gang DP With Red R/R | MK
5 | GI Box 3 x 6 Make | MK
6 | GI Box 3 x 3 Make | MK
7 | GI Box 3 x 3 Deep Make | MK
8 | 6 Gang Grid Switch | MK
9 | 1 Gang 1 Way Switch W/R | MK
10 | 2 Gang 1 Way Switch W/R | MK
11 | 3 Gang 1 Way Switch W/R | MK
12 | Telephone Data outlet | Schneider
13 | RJ 45 Key Stone Jack | Schneider
14 | RJ 45 Jack | Schneider
15 | Patch Panel 24 Port | Schneider
16 | Sub Circuit Monitor (CBS) | Eaton
17 | SVAEL Unit (CBS) | Eaton
18 | Exit Light Wall Mounted | Eaton
19 | Exit Light Hanging Type | Eaton
20 | LED Panel 60 x 60 36W | Ledvance
21 | LED Down Light | Ledvance
22 | LED Spot Light | Ledvance
23 | Card Reader HID RP 40 | HID
24 | Magnetic Lock | Generic
25 | Push Button | Generic
26 | Belden Cable | Belden
27 | Cable Cat6 UTP | Schneider
28 | FP 200 Wire - White & RED | Ducab
29 | Flexible Wire | Ducab
30 | Wire – Single Core | Ducab
31 | GI Trunking Fitting HDG ( Hot Dipped Galvanized) | Unitech-UAE
32 | Earth Strip | Unitech-UAE
33 | PVC Conduit | Decduct
34 | PVC Conduit Accessories | Decduct
35 | PVC Connector | Generic
36 | PVC Conduit fittings | Decoduct
37 | Kopex Gland Complete | Barton
38 | GI Flexible Conduit | Barton
39 | Kopex Gland | Copex
40 | PVC circular Junction Box , Saddel, and Cover | Decoduct

## Table 103
41 | SD/MCP/PROG | GE-USA
42 | ELCB | ABB
43 | MCB | ABB
44 | Cable - XLPE | Ducab
45 | Cable - ECC | Ducab
46 | Cable Gland | Bicon
47 | E- Fiercely | Generic

## Table 104
Approximate Quantities are as follows,

## Table 105
Sr.No. | Description | Quantity | Unit
HVAC System
1 | Supply and installation of Heater battery 300x200 | 3 | Nos.
2 | Relocation of existing of Heater battery 300x200 | 7 | Nos
3 | LTHW Pipe and fitting with insulation | 1 | Lot
4 | Valves Packges for Heater Batteries | 5 | Nos.
5 | Prefabricated Duct with Rockwool insulation | 130 | SQM
BMS Works
1 | Thermostats | 5 | Nos.
2 | Contemporary Controller | 1 | No.
3 | VLD FCU Controller | 5 | Nos.
4 | Duct Temperature sensor TT-D-A | 5 | Nos.
5 | Controle Valve 1/2" | 5 | Nos.
6 | Testing and commissioning | 1 | Lot
7 | Beldon calbe for BMS integration | 1 | Roll
8 | 1.5 mm LSF Cable ( DUCAB) | 2 | Roll
Plumbing and Fire Fighting System
1 | WASHBASIN (55CM HALF PADESTAL )with Automatic Mixer (RAK, and KULDI) | 7 | Nos.
2 | PPR Pipe and fitting for the water Supply ( AQUATHERM) | 1 | Lot
3 | ABS GRY Pipe and fitting for Drainage System ( HEPWORTH) | 1 | Lot
4 | Floor Scanning and Core Cutting | 10 | Nos.

## Table 107
Sr.No. | Description | Quantity | Unit
5 | Water Closet, RAK floor mounted Carla . | 2 | Nos.
6 | Shattaf, Paper holder, Flexible 30cm Pipe , | 2 Each | Nos.
5 | Supply and installation Pendant Sprinkler Head NPT Thread K-5.6, 79C Standard response Bronze finish UL /FM ( RA3415) Make Reliable Model F156 | 10 | Nos.
6 | Testing and Commissioning | 1 | Lot
Electrical and Low Currents Systems
1 | Supply and installation LED Panel 60 x 60 36W including switches (Cool Day Light 50,000 Hr life span with five year warranty (Ledvance) | 52 | Nos.
2 | LED Down Light 18W/8” Cool Day light (Ledvance) | 6 | Nos
3 | Light Switches | 18 | Nos.
4 | Socket Outlet ( Twin) | 44 | Nos.
5 | Socket Outlet ( Single ) | 4 | Nos.
6 | Modification of DBs ( Emergency and Normal ) | 1 | Lot
6 | Twin Telephone Data Outlet ( Schneider) | 14 | Nos
7 | Single Telephone Data Outlet ( Schneider) | 2 | Nos
8 | Data Patch Panel for IDF 24 Port Unloaded | 2 | Nos.
9 | RJ-45 Keystone Jack | 30 | Nos.
10 | Supply, installation with Programming of SCM (JSB/ Eaton) | 15 | Nos
11 | Svael Unit (JSB/ Eaton) | 20 | NOs
12 | Hanging Type Exit Ligth With Diretional Signage (Eaton ) | 2 | Nos
13 | Access Controle System ( Honey well) | 2 | Nos
81 | Wall Mounted Exit Light With Directional Signage ( Eaton ) | 1 | Nos
82 | Supply, installation and Programming of 08 Nos Smoke Detecor With Base and Reprogramming of 24 nos existing smoke detector. | 1 | Lot
83 | GI Trunking & Accessories | 1 | Lot
84 | As Built Drawing for MEP Services | 1 | Lot